Procedures for Verifying and Approving

IEEE 1451.4 TDL Item Registrations

Charles H. Jones

16 September 2004

# Overview

As part of the IEEE 1451.4 implementation, it is possible for manufacturers to request registration of templates and some association Template Description Language (TDL) items.  The registration process involves completing a request form (available on the website) along with providing certain documentation related to the item to be registered.  This document provides details on how to verify and approve the registration request.  Once the items are approved, they must be incorporated into the downloadable files available on the website.  This document also provides details on how to update these files.

Although in an ideal situation the person reviewing these requests would have some technical expertise in the appropriate areas, this will not always be possible and the procedures described here do not assume this technical expertise.  There is, however, an assumption of some basic computer skills.  The reviewer will need to know how to manipulate and view files, use WinZip, and execute programs.  Although not strictly necessary, it may be useful to the reviewer to read the accompanying document IEEE 1451.4 Template and TDL Items Registration Request Process Overview and to skim through clauses 6.1 and 7 of the IEEE 1451.4 standard.

In order to implement the approval procedures, it is necessary to understand that a template is a series of commands beginning with a Template command and ending with an EndTemplate command.

The TDL items that may be registered through this request process are:

- Manufacturer defined templates (see clause 7.2.1).

- Property commands (see clause 7.4 and Annex B.)

- Enumerated types using the Enumerate command (see clause 7.4.5.4 and Annex A.3.)

- Physical units using the Physical_Unit command (see clause 7.4.7 and Annex A.4.)

The biggest aspect to verify is syntax, which can be done by using the syntax checker.  Otherwise, there are some miscellaneous details to check as described in the sections below.

## Templates and Template files

A template provides a description of how to interpret the bits in a TEDS.  A template is written in the Template Description Language (TDL).  Thus, a template is a series of commands that instruct an application program on what a specific set of bits mean.

A template file may contain one or more templates along with a Validation_Keycode.  The TDL.exe syntax checker parses template files.  Template files have a default extension of .tdl. The Validation_Keycode provides a checksum to verify that no errors have been created during transmission of the template file.  The validation keycode.exe program calculates the validation keycode.

## The basic template structure

The basic structure of a template is a Template command followed by a series of commands and ending with an EndTemplate command.  The syntax is as follows:

Template <Manufacturer ID>, <ID number of bits>, <Template ID>,<title>

<commands>

EndTemplate

An example is:

Template 0, 8, 13, “Descriptive Title of Template”

EndTemplate

See clause 7.2.1 for more details.

# Verifying a template

The following subsections describe how to verify each aspect of the template registration request.

## Checking the Validation_Keycode

- Place the following files in the same folder:

  - The program validation keycode.exe.

  - The template file provided in the registration request.

- Double click on validation_keycode.exe.

- When prompted to do so enter the name of the template file and hit enter. (Alternatively, instead of double clicking on the exe and entering the template file name, you can drag the template file and drop it on the exe.)

- While the program is running line numbers are displayed.  If the keycode is correct, the program will state “Keycode CORRECT.”

## Syntax

The syntax can be verified using the syntax checker TDL.exe.  This program is part of the TDL Startup Kit.  To check the syntax:

- Place the following files in the same folder

  - The program file TDL.exe.

  - The template file provided in the registration request.

  - The current version of the file property_list.txt.

- Double click on TDL.exe.

- When prompted to do so enter the name of the template file and hit enter.

- While the program is running line numbers are displayed.  If the template is syntactically correct, the program will complete without error and will state “Syntax is correct.”  If the template contains a syntax error, then the program will provide an error message and abort.

## Description

The TDL provides the ability to add comments using the // or Abstract commands.  Verify that, near the top of the template, there is a description of the template.  Ideally, someone with expertise with the types of transducers applicable to the template should review this.  However, this is optional and the main verification is that the description exists, is in full sentences and is grammatically correct.

For example:

Template 0, 8, 13, “Descriptive Title of Template”

// This is a description of the template.

// Important uses of the template should also be included.

EndTemplate

Or

Template 0, 8, 13, “Descriptive Title of Template”

abstract This is a description of the template.

abstract The types of transducers this template is used with

abstract should be indentified.

EndTemplate

## The Template Command

Recall that the syntax of the Template command is:

Template <Manufacturer ID>, <ID number of bits>, <Template ID>,<title>

The following must be verified in the Template command of the template.

- The Manufacturer ID in the Template command must be a registered Manufacturer ID and must be the Manufacturer ID of the manufacturer submitting the request.  That is, check the Manufacturer ID against the Manufacturer ID list and the registration request form.

- The Template ID in the Template command must be unique among registered templates with the same Manufacturer ID.  This requires checking all other templates registered by this manufacturer to verify that the Template ID is unique.

- The Template command must include a title.  Simply verify that a title is provided in the Template command.

## Template Summary Table

Summary tables for templates provide a breakdown of all the fields defined in the template.  The following table format must be used.  There are examples in Annex A for each of the IEEE standard templates.  Note that the select column allows breakdown of different cases possible through the use of the SELECTCASE command.  The bottom row includes a calculation of the range of how many bits are required for the template.  Note that the syntax checker calculates this information.

| Function | Select | Property/ Command | Description | Access | Bits | Data type (and range) | Units |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ID |  | TEMPLATE | Template ID |  | 8 | Integer (value = 35) |  |
| … | … | … | … | … | … | … | …. |
| Total bits required for TEDS (range): |  |  |  |  | XXX to XXX bits |  |  |

### Summary table functions

A thorough review of the functions in the summary table involves a detailed comparison between the table and the template.  For the purposes of verifying the summary table, a brief comparison should be adequate.  The onus is on the requester to provide a complete and accurate summary.  The main thing to look for is simply length.  A long template should have a fairly long summary and, similarly, a short template will have a short summary.

### Summary table total bits required

The last row of the summary table includes a range of bits required for the TEDS.  This is calculated when running the syntax checker.  Thus, verify that the bit counts in the table match those produced by the syntax checker.

# Verifying a property command

The following subsections describe how to verify each aspect of the property command registration request.

## Syntax

In order to check the syntax of the property commands, create a template file containing the syntactical examples provided in the registration request package.  Then follow the syntax checking instructions provided in section 2.2 above.

For example, if you are provided with this syntactical example:

%test_property,”Quality factor @ f res (Q)”, Cal, 7, ConRelRes, 0.3, 0.03, “”, "V/(m/s²)"

Then create a template file containing

Template 0, 8, 13, “Descriptive Title of Template”

%test_property,”Quality factor @ f res (Q)”, Cal, 7, ConRelRes, 0.3, 0.03, “”, "V/(m/s²)"

EndTemplate

## Description

Ideally, someone with expertise with the types of transducers applicable to the template should review this.  However, this is optional and the main verification is that the description exists, is in full sentences and is grammatically correct.

## Duplication

Check the current list of property names in the file Property_list.txt to verify that there are no duplicate names.  The property name is the first item on the property command with prefix %.  In the above example, the property name is “test_property”.

# Verifying an enumeration or physical unit

The following subsections describe how to verify each aspect of the Enumerate or Physical_Unit command registration request.

## Syntax

Create a template with the Enumerate or Physical_Unit command provided in the registration request.  Then follow the syntax checking instructions provided in section 2.1 above.

For example, given these commands:

ENUMERATE YesNoEnum,"No","Yes"

PHYSICAL_UNIT "%",(1, 0,0, 0,0, 0, 0,0,0,0,100,0)

// Dimensionless Ratio: percent

Create a template file containing:

Template 0, 8, 13, “Descriptive Title of Template”

ENUMERATE YesNoEnum,"No","Yes"

PHYSICAL_UNIT "%",(1, 0,0, 0,0, 0, 0,0,0,0,100,0)

// Dimensionless Ratio: percent

EndTemplate

## Duplication

Check the current list of enumerations or physical units in the files enumerations.tdl or physical_units.tdl to verify that there are no duplicate names.  Specifically, the first field after the Enumerate or Physical_Unit command must be unique for all registered enumerations or physical units.  In the example above, the fields of concern are YesNoEnum and “%”.

# Updating the downloadable files

Once a registration request has been verified and approved, the newly registered item must be incorporated into the files that can be downloaded off the website.  The website allows downloading a zip file that contains all the registered templates.  This file actually contains several files:

- IEEE_standard_templates.tdl - A template file containing all the IEEE standard templates.

- Registered_manufacturer_defined_templates.tdl – A template file containing all the registered manufacturer defined templates.

- Registered_template_summary_tables.doc – A file containing the summary tables provided for each registered manufacturer defined template.

- Property_list.txt – A text file containing a simple list of IEEE standard and registered property tags.

- Registered_property_descriptions.doc – A file containing the descriptions of all registered property commands.

- enumerations.tdl – A text file containing IEEE standard enumerations and registered enumerations.

- physical_units.tdl – A template file containing IEEE standard physical units and registered physical units.

Part of the update process is to change the file revision number.  The standard being used in simply to increment the integer after the R, e.g., R0, R1, R2 …

## Updating template related files

- Cut and paste the newly registered template into the file Registered_manufacturer_defined_templates.tdl.

- Run the syntax checker on the file Registered_manufacturer_defined_templates.tdl to verify the template was inserted properly.

- Update the validation_keycode by following the steps in section 5.4.

- Cut and paste the summary table for the new template into the file Registered_template_summary_tables.doc.

- Update the revision number in the file name for both files Registered_manufacturer_defined_templates.tdl and Registered_template_summary_tables.doc.

- Update the revision number inside the file Registered_template_summary_tables.doc

## Updating property related files

- Enter the newly registered property into the Property_list.txt file.  Each property name must be on a single line and must include the %.

- Cut and paste the description for the newly registered property into the file Registered_property_descriptions.doc.  This should also include an identification of the manufacturer that submitted the property.

- Update the revision number inside the file Registered_property_descriptions.doc.

- Update the revision number in the file name Registered_property_descriptions.doc.

## Updating enumeration or physical units related files

- Cut and paste the newly registered Enumerate or Physical_Unit command into the appropriate file enumerations.tdl or physical_units.tdl.

- Run the syntax checker on enumerations.tdl or physical_units.tdl to verify the command was inserted properly.

- Update the validation_keycode by following the steps in section 5.4.

- Update the revision number inside the appropriate file.

- Update the revision number in the file name.

## Updating the Validation_Keycode in .tdl files

- Place the following files in the same folder:

  - The program validation keycode.exe.

  - The template file to be updated.

- Double click on validation_keycode.exe.

- When prompted to do so enter the name of the template file and hit enter. (Alternatively, instead of double clicking on the exe and entering the template file name, you can drag the template file and drop it on the exe.)

- While the program is running line numbers are displayed.  Since the template files have just been modified, the keycode will not be correct.

- Note the correct keycode calculated by the program and insert this into the Validation_Keycode command at the end of the template file.

- Rerun the validation keycode program to verify the keycode is now correct.