## Slide 1: Medical Seminar
Main Differences between IEC 60601-1:1988 (2nd Edition) and IEC 60601-1:2005 (3rd Edition)
By Grant Schmidbauer
Nemko USA, Inc. San Diego, CA

## Slide 2: Fundamental Change in Standard Title – New Focus
2nd Edition: IEC 60601-1: 1988 Medical Electrical Equipment Part 1: General requirements for safety
3rd Edition: IEC 60601-1: 2005 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance

## Slide 3: Main differences New Definitions
Introduction of Risk Management Process: - ISO 14971 – Medical Devices – Application of Risk Management to Medical Devices
Integration of the following collateral standards: - IEC 60601-1-1: Medical Electrical Systems - IEC 60601-1-4: Programmable Electrical Medical Systems

## Slide 4: Main differences New Definitions cont’d
3.17 Components with high integrity characteristics: Component where one or more characteristics ensure that its function is fault-free in relation to the safety requirements of this standard during the expected service life of the ME equipment in normal use and reasonably foreseeable misuse.
3.27 Essential performance: Performance necessary to achieve freedom from unacceptable risk Note: Essential performance is most easily understood by considering whether its absence or degradation would result in an unacceptable risk.

## Slide 5: Main differences New Definitions cont’d
3.28 Expected service life: Maximum period of useful life as defined by the manufacturer.
3.38 Harm: Physical injury or damage to the health of people or animals, or damage to property or the environment.

## Slide 6: Main differences New Definitions cont’d
3.60 Means of protection (MOP): Means for reducing the risk due to electric shock in accordance with the requirements of this standard Note: Means of protection include insulation, air clearances, creepage distances, impedances, and protective earth connections.
3.58 Means of operator protection (MOOP): Means of protection for reducing the risk due to electric shock to persons other than the patient.
3.59 Means of patient protection (MOPP): Means of protections for reducing the risk due to electric shock to the patient.

## Slide 7: Main differences New Definitions cont’d
3.61 Mechanical hazard: Hazard connected with or produced by physical force.
3.72 Objective evidence: Information which can be proven true, based on facts obtained through observation, measurement, test or other means.

## Slide 8: Main differences New Definitions cont’d
3.75 Oxygen rich environment: Environment in which the concentration of oxygen is: a) greater than 25 % for ambient pressures up to 110kPa; or b) the partial pressure of oxygen is greater than 27.5kPa at ambient pressures exceeding 110kPa.
3.81 Peak working voltage: Highest peak or dc value of a working voltage, including repetitive peak impulses generated in the electrical equipment, but not including external transients.

## Slide 9: Main differences New Definitions cont’d
3.117 Single fault safe: Characteristic of ME Equipment or its parts whereby it remains free of unacceptable risk during its expected service life under single fault conditions.
3.139 Working voltage: Highest voltage to which the insulation or the component under consideration is, or can be, subjected when the electrical equipment is operating under conditions of normal use.

## Slide 10: Main differences New Definitions cont’d
3.129 Touch current: Leakage current flowing from the enclosure or from parts thereof, excluding patient connections, accessible to any operator or patient in normal use, through an external path other than the protective earth conductor, to earth or to another part of the enclosure. Note: The meaning of this term is the same as that of “enclosure leakage current” in the first and second editions of this standard. The term has been changed to align with IEC 60950-1 and to reflect the fact that the measurement now applies also to parts that are normally protectively earthed.

## Slide 11: Main differences General Requirements
4.2 Risk management process for ME Equipment or ME Systems: A risk management process complying with ISO 14971 shall be performed. Compliance is checked by inspection of the risk management file to include: - established a risk management process; - established acceptable levels of risk; and - demonstrated that the residual risk(s) is acceptable (in accordance with the policy for determining acceptable risk).

## Slide 12: Main differences General Requirements cont’d
4.3 Essential performance: The manufacturer shall identify which functions of the ME Equipment and ME Systems are essential performance. Where this standard specifies that essential performance is to be maintained following a particular test, the functions determined above shall be used and compliance shall be checked by inspection, and if necessary, by functional test. Note: Where requirements of this standard refer to essential performance, that essential performance is determined by the manufacturer in accordance with the manufacturer’s policy for risk acceptability.

## Slide 13: Main differences General Requirements cont’d
4.4 Expected service life: The manufacturer shall state the expected service life of the ME Equipment or ME System in the risk management file.

## Slide 14: Main differences General Requirements cont’d
4.6 ME Equipment or ME Systems parts that contact the patient: The risk management process shall include an assessment of whether parts that can come into contact with the patient but fall outside of the definition of applied parts shall be subject to the requirements for applied parts. If the risk management process determines that such parts are subject to the requirements for applied parts, then all the relevant requirements and tests of this standard shall apply, except that 7.2.10 (marking of applied parts) does not apply to such parts.

## Slide 15: Main differences General Requirements cont’d
4.8 Components of ME Equipment: All components shall comply with one of the following: a) the applicable safety requirements of a relevant IEC or ISO standard; b) where there is no relevant IEC or ISO standard, the requirements of this standard have to be applied.

## Slide 16: Main differences General Requirements cont’d

## Slide 17: Main differences General Requirements cont’d
4.9 Use of components with high-integrity characteristics in ME Equipment: A component with high-integrity characteristics shall be used when a fault in a particular component can generate an unacceptable risk. Component with high-integrity characteristics shall be selected and evaluated consistent with their conditions of use and reasonably foreseeable misuse during the expected service life of the ME Equipment.

## Slide 18: Main differences General Requirements for Testing
5.9 Determination of Applied parts and accessible parts: 5.9.1 Applied parts are identified by inspection and by reference to the accompanying documents.

## Slide 19: Main differences Identification, marking and documents
7.1.1 Usability of the identification, markings and documents: The manufacturer shall address in a usability engineering process the risk of poor usability associated with the design of the ME Equipment’s identification, marking and documents.

## Slide 20: Main differences Identification, marking and documents cont’d
7.2.3 Consult accompanying documents:

## Slide 21: Main differences Identification, marking and documents cont’d
7.2.8 Output connectors: Output connectors of ME Equipment intended to deliver power shall be marked with the following information: - rated output voltage; - rated current or power (when applicable); - output frequency (when applicable).
7.5 Safety signs: For the purpose of this clause, markings used to convey a warning, prohibition or mandatory action that mitigates a risk that is not obvious to the operator shall be a safety sign selected from ISO 7010.

## Slide 22: Main differences Identification, marking and documents cont’d
7.9 Accompanying documents: ME Equipment shall be accompanied by documents containing at least the instructions for use and a technical description. The accompanying documents shall be regarded as a part of the ME Equipment. New requirements for instructions for use: - 7.9.2.7 Isolation from the supply mains - 7.9.2.8 Start-up procedure - 7.9.2.10 Messages - 7.9.2.11 Shutdown procedure

## Slide 23: Main differences Protection against electrical hazards
8.3 Classification of applied parts: a) An applied part that is specified in the accompanying documents as suitable for direct cardiac application shall be a Type CF applied part. b) An applied part that includes a patient connection that is intended to deliver electrical energy or an electrophysiological signal to or from the patient shall be a Type BF applied part or Type CF applied part

## Slide 24: Main differences Protection against electrical hazards cont’d
c) An applied part not covered by a) or b) shall be a Type B applied part, Type BF applied part or Type CF applied part. d) For a part that is identified according to 4.6 as needing to be subject to the requirements for an applied part (except for marking), the requirements for a Type B applied part shall apply unless the risk management process identifies a need for the requirements for a Type BF applied part or Type CF applied part to apply.

## Slide 25: Main differences Protection against electrical hazards cont’d
8.5.1 Means of protection: ME Equipment shall have two means of protection to prevent applied parts and other accessible parts from exceeding the limits. One means of protection (MOP) = BI Two means of protection (MOP) = D/RI

## Slide 26: Main differences Protection against electrical hazards cont’d
8.5.1.2 Means of patient protection: - Based on same requirements as IEC 60601-1 2nd edition. - Protective earth connections forming a means of patient protection shall comply with the requirements for protective earthing. - A Y1 capacitor complying with IEC 60384-14 is considered equivalent to one means of patient protection provided that it will pass the dielectric strength test for two means of patient protection. Where two capacitors are used in series, they shall each be rated for the total working voltage across the pair and shall have the same nominal capacitance.

## Slide 27: Main differences Protection against electrical hazards cont’d
8.5.1.3 Means of operator protection: - Based on same requirements as IEC 60950-1. - Protective earth connections forming a means of operator protection shall comply with the requirements for protective earthing. - A Y2 capacitor complying with IEC 60384-14 is considered equivalent to one means of operator protection provided that it will pass the dielectric strength test for one means of operator protection. A Y1 capacitor complying with IEC 60384-14 is considered equivalent to two means of operator protection provided that it will pass the dielectric strength test for two means of operator protection. Where two capacitors are used in series, they shall each be rated for the total working voltage across the pair and shall have the same nominal capacitance.

## Slide 28: Main differences Protection against electrical hazards cont’d
8.5.5.2 Defibrillation protection, Energy reduction test: Defibrillation-proof applied parts or patient connections shall incorporate a means so that the defibrillator energy delivered to a 100Ω load is at least 90 % of the energy delivered to this load with the ME Equipment disconnected.

## Slide 29: Main differences Protection against electrical hazards cont’d
8.6.5 Surface coatings: Conductive elements of ME Equipment that have surface coatings of poorly conducting material such as paint, and between which electrical contact is essential to a protective earth connection, shall have the coatings removed at the point of contact unless an investigation of the joint construction and the manufacturing process has demonstrated that the requirements for impedance and current-carrying capacity are assured without the removal of the surface coating.

## Slide 30: Main differences Protection against electrical hazards cont’d
8.7 Leakage currents and patient auxiliary currents: - Enclosure leakage current changed to Touch current - Earth leakage current limit changed to 5mA in NC and 10mA in SFC, however Touch current effectively limits earth leakage to 500uA in SFC - Concept of total patient leakage current added

## Slide 31: Main differences Protection against electrical hazards cont’d
8.8 Insulation: - Replaced previous concept of testing insulation from Live part (A-xx) and from Applied part (B-xx) to new concept of Means of Protection (MOOP, MOPP) - MOOP based on IEC 60950-1 (Table 6: 1.5/3.0kV) - MOPP based on IEC 60601-1 2nd edition (Table 6: 1.5/4.0kV) - 8.8.2 Distance through solid insulation or use of thin sheet material

## Slide 32: Main differences Protection against electrical hazards cont’d
8.9 Creepage distances and Air clearances: - Replaced previous concept of testing insulation from Live part (A-xx) and from Applied part (B-xx) to new concept of Means of Protection (MOOP, MOPP) - MOOP based on IEC 60950-1 (Table 12 Air clearance 2/4mm, Table 15 Creepage 2.5/5mm) - MOPP based on IEC 60601-1 2nd edition (Table 11 Air clearance 2.5/5mm, Creepage 4/8mm)

## Slide 33: Main differences Protection against electrical hazards cont’d
8.9 Creepage distances and Air clearances: Added new requirements for: - 8.9.1.5 ME Equipment rated for high altitudes - 8.9.1.6 Interpolation - 8.9.1.7 Material Groups classification

## Slide 34: Main differences Protection against electrical hazards cont’d
8.9 Creepage distances and Air clearances: Added new requirements for: - 8.9.1.8 Pollution Degree classification - 8.9.1.9 Overvoltage Category classification - 8.9.1.11 Supply mains overvoltage category

## Slide 35: Main differences Protection against electrical hazards cont’d
8.9 Creepage distances and Air clearances: Added new requirements for: - 8.9.1.12 Secondary circuits - 8.9.3 Spaces filled by insulating compound

## Slide 36: Main differences Protection against mechanical hazards
9.1 Mechanical hazards of ME Equipment: - Introduction of many new mechanical hazards (see table 19)

## Slide 37: Main differences Protection against mechanical hazards cont’d

## Slide 38: Main differences Protection against mechanical hazards cont’d
9.6 Acoustic energy (including infra- and ultrasound) and vibration: ME Equipment shall be designed so that human exposure to acoustic energy and vibration shall not result in an unacceptable risk.

## Slide 39: Main differences Protection against radiation HAZARDS
10.4 Lasers and laser light emitting diodes (LEDs): The relevant requirements of IEC 60825-1:1993 apply. If laser light barriers or similar products are used within equipment, they shall comply with the requirements of IEC 60825-1:1993.

## Slide 40: Main differences Protection against excessive temperatures

## Slide 41: Main differences Protection against excessive temperatures cont’d
11.2.2 ME Equipment and ME Systems used in conjunction with oxygen rich environments: In ME Equipment and ME Systems, the risk of fire in an oxygen rich environment shall be reduced as far as possible under normal condition or single fault conditions.

## Slide 42: Main differences Protection against other hazards
11.3 Constructional requirements for fire-proof enclosures of ME Equipment: This sub-clause provides an alternative means of compliance with selected hazardous situations and fault conditions as identified in 13.1.2. In doing so, the following constructional requirements shall be met or specifically analyzed in the risk management file and if not met, specific justification shall also be given.

## Slide 43: Main differences Accuracy of controls and instruments
Introduction of new requirements for: - 12.1 Accuracy of controls and instruments - 12.2 Usability (IEC 60601-1-6) - 12.3 Alarm systems (IEC 60601-1-8)

## Slide 44: Main differences Protection against hazardous outputs
Introduction of new requirements for: - 12.4.5 Diagnostic or therapeutic radiation - 12.4.6 Diagnostic or therapeutic acoustic pressure

## Slide 45: Main differences PEMS
14 Programmable electrical medical systems (PEMS) Integration of the collateral standard into the standard: - IEC 60601-1-4: Programmable Electrical Medical Systems

## Slide 46: Main differences Mechanical strength
15.3 Mechanical strength - Introduction of many new mechanical hazards (see table 28)

## Slide 47: Main differences Mechanical strength cont’d

## Slide 48: Main differences
15.4.3.4 Lithium batteries: Lithium batteries used in ME Equipment that could become a hazard shall comply with the requirements of IEC 60086-4 – Part 4: Safety of lithium batteries.

## Slide 49: Main differences ME Systems
16 Requirements for ME Systems Integration of the collateral standard into the standard: - IEC 60601-1-1: Medical Electrical Systems

## Slide 50: Main differences Guide to marking and labeling
Annex C: Guide to marking and labeling requirements for ME Equipment and ME Systems. - Introduction of various tables for markings and labeling (see tables next pages)

## Slide 51: Main differences Guide to marking and labeling cont’d

## Slide 52: Main differences Guide to marking and labeling cont’d

## Slide 53: Main differences Guide to marking and labeling cont’d

## Slide 54: Main differences Guide to marking and labeling cont’d

## Slide 55: Main differences Guide to marking and labeling cont’d

## Slide 56: Contact Information
Grant Schmidbauer General Manager Nemko USA, Inc. 11696 Sorrento Valley Road San Diego, CA 92121 USA Tel: 858-755-5525 Ex. 208 Fax: 858-452-1810 Email: grant.schmidbauer@nemko.com Website: www.nemko.com