IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs) |  |
| --- | --- | --- |
| Title | Channel Model for Body Area Network (BAN) |  |
| Date Submitted | [14 May, 2008] |  |
| Source | [Kamya Yekeh Yazdandoost] Medical ICT Institute, NICT New Generation Wireless Communication research Center, 3-4 Hikarino-oka Yokosuka 239-0847, Japan [Kamran Sayrafian-Pour] Information Technology Laboratory National Institute of Standard & Technology Gaithersburg, MD 20899 USA | Voice : +81-45-847-5435 Fax : +81-45-847-5431 E-mail: [yazdandoost@nict.go.jp] Voice : +1-301-975-5479 E-mail :[ksayrafian@nist.gov] |
| Re: | [Body Area Network (BAN) Channel Model document] |  |
| Abstract | [This is a draft document of the IEEE802.15.6 channel modeling subcommittee. It provides how channel model should be developed for body area network. |  |
| Purpose | [The purpose of this document is to provide the work of the channel modeling subcommittee and recommendations on how the channel model for BAN can be used. |  |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |  |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |  |

Channel Modeling Subcommittee Report

| Date | Revision No. |
| --- | --- |
| 11/15/2007 | 15-07-0943-00-0ban |
| 01/11/2008 | 15-08-0033-00-0006 |

Table of Contents

# 1. Introduction………………….………………………………………………………………3

# 2. Definitions & Overview..……….………………………………………………………….3

# 3. Scenarios……………………………………………………………………………………3

# 4. Antenna effect ……………………………………………………………………………..4

# 5. Electrical properties of body tissues……………………………………………………6

# 6. Channel characterization ………………………………………………………………...6

6.1. Model types.………....………………………..………………………………………….6

6.2. Path loss…………………….…………………………………………………………….7

6.3. Shadowing… ………………………..……………..……………………………………..8

# 7. List of contributors……....…………………………………………………………………8

# 8. References ……….……....…………………………………………………………………8

# 1. Introduction

This is a document for the IEEE802.15.6 (Body Area Network) channel modeling subcommittee. This document provides recommendations of the channel modeling subcommittee of IEEE802.15.6. The models discussed generally characterize the path loss of BAN devices taking into account possible shadowing due to the human body or obstacles near the human body and postures of human body.

The channel model is needed to evaluate the performance of different physical layer proposals. The main goal of these channel models is a fair comparison of different proposals. They are not intended to provide information of absolute performance in different environments or body postures. The number of available measurements on which the model can be based, in the frequency range of …….and ……

# 2. Definitions & Overview

An important step in the development of a wireless body area network is the characterization of the electromagnetic wave propagation from devices that are close to or inside the human body. The complexity of the human tissues structure and body shape make it difficult to drive a simple path loss model for BAN. As the antennas for BAN applications are placed on or inside the body, the BAN channel model needs to take into account the influence of the body on the radio propagation.

For the purpose of this document, we define 3 types of nodes as follows:

- Implant node: A node that is placed inside the human body. This could be immediately below the skin to further deeper inside the body tissue

- Body Surface node: A node that is placed on the surface of the human skin or at most 2 centimeters away

- External node: A node that is not in contact with human skin (between a few centimeters and up to 5 meters away from the body)

For body surface communication, the distance between the transmitting and receiving nodes shall consider the distance around the body if transmitter and receiver are not placed in the same side rather than straight line through the body. This allows creeping wave diffraction to be also considered. For external node communication, the distance between transmitter and receiver shall be from the body vicinity or inside body to 2 meters away. In some cases, the maximum range for medical device shall be 5 meters.

The maximum power limitation for on-body medical device shall be TBD.

The maximum power limitation for MICS is [1], [2]:

- ETSI (European Telecommunications Standards Institute): The output power is set to a maximum of 25 uW ERP.

- FCC & ITU-R: The output power is set to a maximum of 25 uW EIRP, which is ≈ 2.2 dB lower than the ERP level.

- The 25 uW limit applies to the signal level outside of the body (total radiating system), which allows for implant power levels to be increased to compensate for body losses.

Frequency band for implant devices (i.e. MICS) shall be 402-405 MHz as specified in [8]. Frequency band for other devices are TBD.

The structure of the channel model for scenarios involving body surface and implant is not similar. The channel model for implant device is fundamentally different.

# 3. Scenarios

From [9,6], a list of scenarios can be identified in which IEEE802.15.6 devices will be operating. These scenarios along with their description and frequency band are listed in Table 1. The scenarios are determined based on the location of the communicating nodes (i.e. implant, body surface and external). The scenarios are grouped into classes that can be represented by the same Channel Models (CM).

| Scenario | Description | Frequency Band | Channel Model |
| --- | --- | --- | --- |
| S1 | Implant to Implant | 402-405 MHz | CM1 |
| S2 | Implant to Body Surface | 402-405 MHz | CM2 |
| S3 | Implant to External | 402-405 MHz | CM2 |
| S4 | Body Surface to Body Surface (LOS) | TBD (f1,… fn) | CM3 |
| S5 | Body Surface to Body Surface (NLOS) | TBD (f1,… fn) | CM3 |
| S6 | Body Surface to External (LOS) | TBD (f1,… fn) | CM4 |
| S7 | Body Surface to External (NLOS) | TBD (f1,… fn) | CM4 |

Table 1: List of scenarios and their descriptions

The distance of external devices is considered to be a maximum of 5 meters.

Possible channel models described above are graphically displayed in Fig. 1.

Fig. 1: Possible communication links for Body Area Networking

# 4. Antenna Effect

An antenna placed on the surface or inside a body will be heavily influenced by its surroundings [3].  The consequent changes in antenna pattern and other characteristics needs to be understood and accounted for during any propagation measurement campaign.

The form factor of an antenna will be highly dependent on the requirements of the application. For MICS applications, for example, a circular antenna may be suitable for a pacemaker implant, while a helix antenna may be required for a stent or urinary implant.  The form factor will affect the performance of the antenna and, the antenna performance will be very important to the overall system performance. Therefore, an antenna which has been designed with respect to the body tissues (or considered the effect of human body) shall be used for the channel model measurements [4].

The BAN antennas may be classified into two main groups [5]:

- Electrical antennas, such as dipole

Electrical antenna- typically generates large components of E-field normal to the tissues interface, which overheat the fat tissue. This is because boundary conditions require the normal E-field at the interface to be discontinuous by the ratio of the permittivities, and since fat has a lower permittivity than muscle, the E-field in the fat tissue is higher.

- Magnetic antennas, such as loop

Magnetic antenna produces an E-field mostly tangential to the tissues, which seem not to couple as strongly to the body as electrical antennas. Therefore, does not overheat the fat.

There are antennas same as helical-coil, which is similar to a magnetic antenna in some respect, but its heating characteristics appear to be more like an electrical antenna. The strong E-field generated between the turns of coil is mainly responsible for tissue heating.

It should be noted that SAR in the near field of the transmitting antenna depends mainly on the H-field; however, SAR in the far field of the transmitting antenna depends mainly on the E-field.

# 5. Electrical properties of body tissues

The human body is not an ideal medium for radio frequency wave transmission. It is partially conductive and consists of materials of different dielectric constants, thickness, and characteristic impedance. Therefore depending on the frequency of operation, the human body can lead to high losses caused by power absorption, central frequency shift, and radiation pattern destruction. The absorption effects vary in magnitude with both frequency of applied field and the characteristics of the tissue [10, 11, 12, 13].

# 6. Channel characterization

# 6.1. Model types

In all cases, two types of model may be generated:

- A theoretical or mathematical model

- An empirical model

A theoretical model may be traceable back to first principles and will permit precise modeling of a specific situation at radio link level.  It is intended for detailed exploration of, for example, the influence of body structures on antenna patterns.  It will require a detailed description of the propagation environment and is therefore probably not suitable for modeling of macro environments.

An empirical model may be traceable to an agreed set of propagation measurements and is intended to provide a convenient basis for statistical modeling of networks.  Compared to the theoretical model, the empirical model will use a greatly simplified description of the environment and, although statistically accurate at network level, will not be precise at link level.

Appropriate efforts will be made to ensure that the two sets of models are consistent with each other.

# 6.2. Path loss

Unlike traditional wireless communications, the path loss for body area network system (on body applications), is both distance and frequency dependent. The frequency dependence of body tissues shall be considered.

The path loss model in dB between the transmitting and the receiving antenna as a function of the distance d based on the Friis formula in free space is described by [14, 15]:

PL(d)=PL0+10nlog10dd0                                                  (1)

where PL0 is the path loss at a reference distance d0 which is set to TBD, and n is the path-loss exponent, TBD.

The reference path loss near the antenna depends on the separation between the antenna and the body due to antenna mismatch. This mismatch indicates that a body-aware antenna design could improve system performance.

The frequency dependent path loss is given by [7]:

PL(f)=PrfPTXf=12PL0ηTXfηTRfffc−2k+1dd0n                        (2)

where Pr is received power, PTX is transmitted power, ηTX and ηRX are the efficiency of transmitting and receiving antennas, fc is the central frequency, and k is the frequency dependence coefficients.

In case of MICS, for only inside the body cavity, it is assumed that distance dependent path loss is negligible if device placed near to the body surface and only frequency dependent path loss with respect to the type of tissues (amount of absorption) shall be modeled.

6.3. Shadowing

Due to the variation in the surrounding of human body or even movement of body parts, the received power will be different from the mean value for a given distance as shown in equation (1). This phenomenon is called shadowing, reflects the path loss variation around the mean. The shadowing should be considered for stationary position of human as well as for the body movements.

When considering shadowing, the total path loss PL can be expressed by:

PL=PL(d)+S                                                   (3)

where PL(d) is expressed by the equation (1) and S represents the shadowing component.

# 7. List of contributors

Arthur Astrin

Rob J Davise

# 8. References

- ERC Recommendation 70-03 relating to the use of Short Range Device (SRD), European Conference of Postal and Telecommunications Administrations, CEPT/ERC 70-03, Tromsø, Norway, 1997.

- FCC, Medical implant communications, January 2003,

http://wireless.fcc.gov/services/index.htm?job=service_home&id=medical_implant

- W.-T. Chen; H.-R. Chuang, “Numerical computation of human interaction with arbitrarily oriented superquadric loop antennas in personal communications,” IEEE Trans. on Antenna and Propagation, vol.46, no. 6, pp. 821-828, June 1998.

- Kamya Y. Yazdandoost and Ryuji Kohno, “The Effect of Human Body on UWB BAN Antennas,” IEEE802.15-07-0546-00-0ban.

- Kamya Y. Yazdandoost and Ryuji Kohno, “Wireless Communications for Body Implanted Medical Device,” Asia Pacific Microwave Conference, APMC2007, pp.

- Kamya Y. Yazdandoost et al, “Channel Characterization for BAN Communications,” IEEE802.15-07-0641-00-0ban.

- Andreas F. Molisch et al, “A Comprehensive Model for Ultrawideband Propagation Channels,” IEEE Global Telecommunications Conference, GLOBECOM '05. Vol.6, pp. 3648-3653.

- 15-07-0939-01-0ban-ieee-802-15-6-regulation-subcommittee-report

- 15-07-0735-06-0ban-ban-application-matrix_amaledit

- C. H. Duney, H. Massoudi, and M. F. Iskander, “Radiofrequency radiation dosimetry handbook,” USAF School of Aerospace Medicine, October 1986.

- C. Gabriel and S. Gabriel, “Compilation of the dielectric properties of body tissues at RF and microwave frequencies,” AL/OE-TR-1996-0037, June 1996, http://www.brooks.af.mil/AFRL/HED/hedr/reports/dielectricReport/Report.html.

- Italian National Research Council, Institute for Applied Physics, “Dielectric properties of body tissues,” http://niremf.ifac.cnr.it

- P. Gandhi, “.Biological Effects and Medical Applications of Electromagnetic Energy,” Prentice Hall, Englewood Cliffs, N.J., 1990.

- E. Reusens, W. Joseph, G. Vermeeren, and L. Martens, „On-body measurements and characterization of wireless communication channel fro arm and torso of human,“ International Workshop on Wearable and Implantabel Body Sensor Networks, BSN07, Achen, March 2007, pp. 26-28.

- A. Fort, J. Ryckaert, C. Desset, P. De Doncker, P. Wambacq, and L. Van Biesen, « Ultra-wideband channel model for communication around the human body, ” IEEE Journal on Selected Areas in Communications, vol. 24, pp.927-933, April 2006.