IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs) |  |
| --- | --- | --- |
| Title | Channel Model for Body Area Network (BAN) |  |
| Date Submitted | [8 September, 2008] |  |
| Source | [Kamya Yekeh Yazdandoost] Medical ICT Institute, NICT New Generation Wireless Communication research Center, 3-4 Hikarino-oka Yokosuka 239-0847, Japan [Kamran Sayrafian-Pour] Information Technology Laboratory National Institute of Standard & Technology Gaithersburg, MD 20899 USA | Voice : +81-45-847-5435 Fax : +81-45-847-5431 E-mail: [yazdandoost@nict.go.jp] Voice : +1-301-975-5479 E-mail :[ksayrafian@nist.gov] |
| Re: | [Body Area Network (BAN) Channel Model document] |  |
| Abstract | [This is a draft document of the IEEE802.15.6 channel modeling subcommittee. It provides how channel model should be developed for body area network. |  |
| Purpose | [The purpose of this document is to provide the work of the channel modeling subcommittee and recommendations on how the channel model for BAN can be used. |  |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |  |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |  |

Channel Modeling Subcommittee Report

| Date | Revision No. |  |
| --- | --- | --- |
| 11/15/2007 | 15-07-0943-00-0ban |  |
| 01/11/2008 | 15-08-0033-00-0006 |  |
| 05/14/2008 | 15-08-0033-01-0006 |  |
| 05/15/2008 | 15-08-0033-02-0006 |  |
| 07/16/2008 | 15-08-0033-03-0006 |  |

Table of Contents

# 1. List of Contributors……………………………………………………...…………………5

# 2. Introduction………………….………………………………………………………………6

# 3. Definitions & Overview..……….………………………………………………………….7

# 4. Scenarios…………………………………………………………………………………….8

# 5. Antenna Effect ……………………………………………………………………………...9

# 6. Electrical Properties of Body Tissues…..…………………………………………….10

# 7. Channel Characterization ………………………………………….………...………….10

7.1. Model types.………....………………………..…………………….…………………...10

7.2. Path loss…………………….…………………………………………………………...11

7.3. Shadowing… ………………………..……………..……………………………………11

7.4. Power delay profile.………………..……………..…………………………………..…11

7.5. Fading… ………………………..……………..………………………………………..12

7.5.1. Small scale fading………..……………..………….………….………………….12

7.5.2. Large scale fading………..……………..………….………….………………….12

# 8. Models and Scenarios…….………………………………………….………...………..12

8.1. In-Body…….…………….......………………………..…….……….…………………12

8.1.1. Implant to Implant CM1 (Scenario S1) for 402-405 MHz ……………………...13

8.1.2. Implant to body surface CM2 (Scenario S2 ) for 402-405 MHz ………………..13

8.1.3. Implant to external CM2 (Scenario S3) for 402-405 MHz…………….………..13

8.2. On-Body………………………………………………………………………………...14

8.2.1. Body surface to body surface CM3 (Scenarios S4 & S5) for 13.5 MHz………..14

8.2.2. Body surface to body surface CM3 (Scenario S4 & S5) for 5-50 MHz ………..15

8.2.3. Body surface to body surface CM3 (Scenario S4 & S5) for 400 MHz…………16

8.2.4. Body surface to body surface CM3 (Scenario S4 & S5) for 600 MHz…………16

8.2.5. Body surface to body surface CM3 (Scenario S4 & S5) for 900 MHz…………17

8.2.5.A……………………………………………………………………………..17

8.2.5.B……………………………………………………………………………..18

8.2.5.C……………………………………………………………………………..19

8.2.6. Body surface to body surface CM3 (Scenario S4 & S5) for 2.4 GHz..…………19

8.2.6. A…………………………………………………………………………….19

8.2.6.B……………………………………………………………………………..20

8.2.6.C……………………………………………………………………………..21

8.2.7. Body surface to body surface CM3 (Scenario S4 & S5) for 3.1-10.6 GHz…….22

8.2.7.A…………………………………………………………………………….22

8.2.7.B…………………………………………………………………………….22

8.2.8. Body surface to external CM4 (Scenario S6 & S7) for 900 MHz…………23

8.2.9. Body surface to external CM4 (Scenario S6 & S7) for 2.4 GHz…..………23

8.2.10. Body surface to external CM4 (Scenario S6 & S7) for 3.1-10.6 GHz..….23

# 9. References ….……..…...…………………..…………………………….……………….24

# 1. List of contributors

Major contributions were received from the following individuals:

Arthur Astrin

Takahiro Aoyagi

Rob J Davise

Guido Dolmans

Andrew Fort

Ban Gilbert

Leif Hanlen

Jung-Hwan Hwang

John Hagedorn

Marco Hernandez

Norihiko Katayama

Takehiko Kobayashi

Ryuji Kohno

Sung-Weon Kang

Huan-bang Li

Daniel Lewis

Dino Miniutti

Il-Hyoung Park

David Rodda

Kamran Sayrafian-Pour

David Smith

Jun-ichi Takada

Kenichi Takizawa

Judith Terrill

Kamya Yekeh Yazdandoost

Wenbin Yang

Andrew Zhang

# 2. Introduction

This document summarizes the activities and recommendations of the channel modeling subgroup of IEEE802.15.6 (Body Area Network). The Task Group TG6 is intended to develop Body Area Network for medical and non-medical devices that could be placed inside or on the surface of human body.

The models discussed generally characterize the path loss of BAN devices taking into account possible shadowing due to the human body or obstacles near the human body and postures of human body.

The channel model is needed to evaluate the performance of different physical layer proposals. The main goal of these channel models is a fair comparison of different proposals. They are not intended to provide information of absolute performance in different environments or body postures. The list of frequency band and number of available measurements on which the model can be based is shown in Table 1.

| Description |  | Frequency Band |
| --- | --- | --- |
| Implant |  | 402-405 |
| On-Body |  | 13.5 MHz |
| On-Body |  | 5-50 MHz (HBC) |
| On-Body |  | 400 MHz |
| On-Body | 600 MHz |  |
| On-Body | 900 MHz |  |
| On-Body | 2.4 GHz |  |
| On-Body | 3.1-10.6 GHz |  |

Table 1: List of frequency band

Since the subgroup was formed, a large number of documents has been submitted to the channel modeling subgroup or presented and discussed at IEEE802.15.6 meetings and teleconference calls. They can be found on the https://mentor.ieee.org/802.15/documents, and are cited where appropriate in this document. The channel model subgroup started its activities at the meeting in (Date and Location), and is submitting this final report in November 2008 (Dallas). Appreciation is extended to all the participants from academia and industry, whose efforts made this model possible.

Despite of significant efforts have been carried out to make models as realistic as possible, the number of available measurements on which the model can be based for wide range of frequencies (see Table 1) are insufficient.

To facilitate the use of the model, this document also includes a MATLAB program for the generation of each channel model.

The reminder of the document is organized in the following way: section 3 gives an overview as well as definition. Section 4 describes different scenarios and range of frequencies. Section 5 gives a short discussion of antenna. Section 6 gives an overview of medium. Section 7 provides channel characterization.  Section 8 gives full detail on models and scenarios.

# 3. Definitions & Overview

An important step in the development of a wireless body area network is the characterization of the electromagnetic wave propagation from devices that are close to or inside the human body. The complexity of the human tissues structure and body shape make it difficult to drive a simple path loss model for BAN. As the antennas for BAN applications are placed on or inside the body, the BAN channel model needs to take into account the influence of the body on the radio propagation.

For the purpose of this document, we define 3 types of nodes as follows:

- Implant node: A node that is placed inside the human body. This could be immediately below the skin to further deeper inside the body tissue

- Body Surface node: A node that is placed on the surface of the human skin or at most 2 centimeters away

- External node: A node that is not in contact with human skin (between a few centimeters and up to 5 meters away from the body)

For body surface communication, the distance between the transmitting and receiving nodes shall consider the distance around the body if transmitter and receiver are not placed in the same side rather than straight line through the body. This allows creeping wave diffraction to be also considered. For external node communication, the distance between transmitter and receiver shall be from the body vicinity or inside body to 2 meters away. In some cases, the maximum range for medical device shall be 5 meters.

The maximum power limitation for on-body medical device shall be TBD.

The maximum power limitation for MICS is [1], [2]:

- ETSI (European Telecommunications Standards Institute): The output power is set to a maximum of 25 uW ERP.

- FCC & ITU-R: The output power is set to a maximum of 25 uW EIRP, which is ≈ 2.2 dB lower than the ERP level.

- The 25 uW limit applies to the signal level outside of the body (total radiating system), which allows for implant power levels to be increased to compensate for body losses.

Frequency band for implant devices (i.e. MICS) shall be 402-405 MHz as specified in [8].

The structure of the channel model for scenarios involving body surface and implant is not similar. The channel model for implant device is fundamentally different.

# 4. Scenarios

From [9,6], a list of scenarios can be identified in which IEEE802.15.6 devices will be operating. These scenarios along with their description and frequency band are listed in Table 2. The scenarios are determined based on the location of the communicating nodes (i.e. implant, body surface and external). The scenarios are grouped into classes that can be represented by the same Channel Models (CM).

| Scenario | Description | Frequency Band | Channel Model |
| --- | --- | --- | --- |
| S1 | Implant to Implant | 402-405 MHz | CM1 |
| S2 | Implant to Body Surface | 402-405 MHz | CM2 |
| S3 | Implant to External | 402-405 MHz | CM2 |
| S4 | Body Surface to Body Surface (LOS) | 13.5, 50, 400, 600, 900 MHz 2.4, 3.1-10.6 GHZ | CM3 |
| S5 | Body Surface to Body Surface (NLOS) | 13.5, 50, 400, 600, 900 MHz 2.4, 3.1-10.6 GHZ | CM3 |
| S6 | Body Surface to External (LOS) | 13.5, 50, 400, 600, 900 MHz 2.4, 3.1-10.6 GHZ | CM4 |
| S7 | Body Surface to External (NLOS) | 13.5, 50, 400, 600, 900 MHz 2.4, 3.1-10.6 GHZ | CM4 |

Table 2: List of scenarios and their descriptions

The distance of external devices is considered to be a maximum of 5 meters. Possible channel models described above are graphically displayed in Fig. 1.

Fig. 1: Possible communication links for Body Area Networking

# 5. Antenna Effect

An antenna placed on the surface or inside a body will be heavily influenced by its surroundings [3].  The consequent changes in antenna pattern and other characteristics needs to be understood and accounted for during any propagation measurement campaign.

The form factor of an antenna will be highly dependent on the requirements of the application. For MICS applications, for example, a circular antenna may be suitable for a pacemaker implant, while a helix antenna may be required for a stent or urinary implant.  The form factor will affect the performance of the antenna and, the antenna performance will be very important to the overall system performance. Therefore, an antenna which has been designed with respect to the body tissues (or considered the effect of human body) shall be used for the channel model measurements [4].

The BAN antennas may be classified into two main groups [5]:

- Electrical antennas, such as dipole

Electrical antenna- typically generates large components of E-field normal to the tissues interface, which overheat the fat tissue. This is because boundary conditions require the normal E-field at the interface to be discontinuous by the ratio of the permittivities, and since fat has a lower permittivity than muscle, the E-field in the fat tissue is higher.

- Magnetic antennas, such as loop

Magnetic antenna produces an E-field mostly tangential to the tissues, which seem not to couple as strongly to the body as electrical antennas. Therefore, does not overheat the fat.

There are antennas same as helical-coil, which is similar to a magnetic antenna in some respect, but its heating characteristics appear to be more like an electrical antenna. The strong E-field generated between the turns of coil is mainly responsible for tissue heating.

It should be noted that SAR in the near field of the transmitting antenna depends mainly on the H-field; however, SAR in the far field of the transmitting antenna depends mainly on the E-field.

# 6. Electrical Properties of Body Tissues

The human body is not an ideal medium for radio frequency wave transmission. It is partially conductive and consists of materials of different dielectric constants, thickness, and characteristic impedance. Therefore depending on the frequency of operation, the human body can lead to high losses caused by power absorption, central frequency shift, and radiation pattern destruction. The absorption effects vary in magnitude with both frequency of applied field and the characteristics of the tissue [10, 11, 12, 13].

# 7. Channel Characterization

# 7.1. Model Types

In all cases, two types of model may be generated:

- A theoretical or mathematical model

- An empirical model

A theoretical model may be traceable back to first principles and will permit precise modeling of a specific situation at radio link level.  It is intended for detailed exploration of, for example, the influence of body structures on antenna patterns.  It will require a detailed description of the propagation environment and is therefore probably not suitable for modeling of macro environments.

An empirical model may be traceable to an agreed set of propagation measurements and is intended to provide a convenient basis for statistical modeling of networks.  Compared to the theoretical model, the empirical model will use a greatly simplified description of the environment and, although statistically accurate at network level, will not be precise at link level.

Appropriate efforts will be made to ensure that the two sets of models are consistent with each other.

# 7.2. Path Loss

Unlike traditional wireless communications, the path loss for body area network system (on body applications), is both distance and frequency dependent. The frequency dependence of body tissues shall be considered.

The path loss model in dB between the transmitting and the receiving antenna as a function of the distance d based on the Friis formula in free space is described by [14, 15]:

PL(d)=PL0+10nlog10dd0                                                  (1)

where PL0 is the path loss at a reference distance d0 which is set to TBD, and n is the path-loss exponent, TBD.

The reference path loss near the antenna depends on the separation between the antenna and the body due to antenna mismatch. This mismatch indicates that a body-aware antenna design could improve system performance.

# 7.3. Shadowing

Due to the variation in the surrounding of human body or even movement of body parts, the received power will be different from the mean value for a given distance as shown in equation (1). This phenomenon is called shadowing, reflects the path loss variation around the mean. The shadowing should be considered for stationary position of human as well as for the body movements.

When considering shadowing, the total path loss PL can be expressed by:

PL=PL(d)+S                                                   (3)

where PL(d) is expressed by the equation (1) and S represents the shadowing component.

# 7.4. Power Delay Profile

Because of multipath reflections, the channel response of a wireless body area network channel looks likes a series of pulses. In practice the number of pulses that can be distinguished is very large, and depends on the time resolution of the measurement system. The power delay profile of channel is an average power of the channel as a function of an excess delay with respect to the first arrival path.

# 7.5. Fading

Channel Fading

Large Scale Fading

Small Scale Fading

TBA

# 7.5.1. Small Scale Fading

Small scale fading refers to the rapid changes of the amplitude and phase of the received signals within a small local area due to small changes in location of the on body device or body positions, in a given short period of time. The small scale fading based on multipath time delay spread can be divided to the flat fading and frequency selective fading.

# 7.5.2 Large Scale Fading

Large scale fading refers to the fading due to motion over large areas; this is referring to the distance between antenna position on the body and at location in the room (home, office, or hospital).

Averaging the attenuation between each antenna position on the body and each antenna location in the room will removes the effect of small scale fading due to small changes in the user position around the room.

# 8. Models and Scenarios

# 8.1. In-Body

Since physical measurement and experimental study inside human body is not feasible, a 3D simulation & visualization scheme was used to study the propagation characteristics of MICS. The human body model used in this study includes (frequency dependent) dielectric properties of 300+ parts in a male human body with a maximum resolution of 2mm. The implant antenna used in this study is a multi-thread loop antenna with the following characteristics [17]:

- Size: 8.2 x 8.1 x 1 mm

- Metallic Layer: Copper, t=0.036 mm

- Substrate: D51 (NTK), εr=30,tan θ = 0.000038, and t=1mm

- The metallic layer is covered by RH-5, εr=1.0006, tan θ= 0, t =1mm

Parameters of a statistical path loss model have been extracted that fits the following equation.

PL(d) = PL(d0) + 10nlog10(dd0) + S  where  S ~ N(0, σs) and d0 = 50mm

The parameters corresponding to CM1 and CM2 are expressed in the tables in section 8.1.1 and 8.1.2. Details of the model derivation can be found in [19,20].

# 8.1.1. Implant to Implant CM1 (Scenario S1) for 402-405 MHz

| Implant to Implant | PL(d0)(dB) | n | σs(dB) |
| --- | --- | --- | --- |
| Deep Tissue | 35.04 | 6.26 | 8.18 |
| Near Surface | 40.94 | 4.99 | 9.05 |

# 8.1.2. Implant to body surface CM2 (Scenario S2 ) for 402-405 MHz

| Implant to Body Surface | PL(d0)(dB) | n | σs(dB) |
| --- | --- | --- | --- |
| Deep Tissue | 47.14 | 4.26 | 7.85 |
| Near Surface | 49.81 | 4.22 | 6.81 |

Remark: A distance of up to 20mm directly from the body surface has been considered in the derivation of the above parameters. One should keep in mind that layers of clothing could cause additional loss to the signal.

# 8.1.3. Implant to external CM2 (Scenario S3) for 402-405 MHz

The scenario S3 for CM2 can be approximated by considering a combination of scenarios S2 and S6 (or S7). In the simple case of an environment where there are no objects or obstacles, a free space path loss can be added to CM2 to account for the additional loss that the implant signal will go through once it leaves the body. This usually occurs after around 10 cm away from the body surface.

On the other hand, if there are objects within the 5 meter distance of the human body, a channel model for scenarios S6 or S7 (at 400 MHz) can account for the impact of these objects; and therefore, such model can be added to CM2 to emulate scenario S3. (Remark: the path loss model for S6 (or S7) in this case should not include the on-body TX antenna gain pattern.)

# 8.2. On-body

# 8.2.1. Body surface to body surface CM3 (Scenarios S4 & S5) for 13.5 MHz

Measurement for frequency range of 13.550-13.571 MHz has been performed in [21]. The important observation at these frequencies is that the body channel exhibits path loss that is nearly similar to free space. However, the available bandwidth is quite small (21 KHz). The following table summarizes the measurement results. Further detail on set-up, derivation and data analysis can also be found in [21].

| Description | Signal amplitude reduction |  |  |  |  |  |  | dB loss in relation to air |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Through the hand | 3.3% |  |  |  |  |  |  | -0.15 |  |  |  |  |
| Through the wrist | 2.8% |  |  |  |  |  |  | -0.12 |  |  |  |  |
| Torso, front to back | 3.4% |  |  |  |  |  |  | -0.15 |  |  |  |  |
| Through the thigh | 1.9% |  |  |  |  |  |  | -0.08 |  |  |  |  |
| Through the ankle |  | 2.8% |  |  |  |  |  |  | -0.12 |  |  |  |
| Left ear to right ear |  | 2.0% |  |  |  |  |  |  | -0.09 |  |  |  |
| Left ear to right ear, wearing metal glasses |  | 1.5% |  |  |  |  |  |  | -0.07 |  |  |  |
| Distance through body (cm) |  | 7.6 | 10.2 | 12.7 | 15.2 | 17.8 | 20.3 |  | 22.9 | 25.4 | 27.9 | 30.5 |
| Received signal P-P (mV) |  | 6.93 | 4.74 | 2.93 | 1.86 | 1.25 | 0.837 |  | 0.633 | 0.471 | 0.384 | 0.305 |

# 8.2.2. Body surface to body surface CM3 (Scenario S4 & S5) for 5-50 MHz

Human body can also be used as a communication media over the range of frequencies 5-50MHz. No modulation is needed in this form of communication which is referred to as Human Body Communication (HBC). The channel model for HBC is composed of the frequency response and the noise characteristics. Measurement process to obtain this frequency response has been outlined in [22].

The following table and transfer function summarizes the result of this measurement that corresponds to fingertips of one hand to the other. Since, this scenario causes the most amplitude attenuation, these results have been chosen for the HBC channel model.

| Parameters |  | Conditions |
| --- | --- | --- |
| Locations of transmitter and receiver |  | Fingertips of each hand |
| Transmission distance (cm) |  | 150 |
| Contact location of signal electrode |  | Fingertip of thumb |
| Size of signal electrode (cm2) | 2x2 |  |
| Load impedance of receiver | 10 MΩ |  |

# 8.2.3. Body surface to body surface CM3 (Scenario S4 & S5) for 400 MHz

The following path loss model is based on measurements that cover frequencies of 400-450 MHz. Further measurement set up, derivation and data analysis can be found in [23]. The table below summarizes the model and corresponding parameters.

|  |  | Hospital Room |  | Anechoic Chamber |
| --- | --- | --- | --- | --- |
| Path loss model |  | PLd [dB]=a⋅log10d+b+c+N |  |  |
| a |  | 20.6 |  | 46.4 |
| b |  | 12.4 |  | -40.4 |
| c |  | -16.5 |  | -16.5 |
| σN | 7.2 |  | 2.7 |  |

- a and b : Coefficients of linear fitting

- d : Tx-Rx distance in mm.

- N : Normally distributed variable with standard deviation N

- c:  The difference between the averaged signal levels of the dipole or colinear antenna used in the whole measurement and that of the chip antenna

# 8.2.4. Body surface to body surface CM3 (Scenario S4 & S5) for 600 MHz

The following path loss model is based on measurements that cover frequencies of 608-614 MHz. Further measurement set up, derivation and data analysis can be found in [23]. The table below summarizes the model and corresponding parameters.

|  |  | Hospital Room |  | Anechoic Chamber |
| --- | --- | --- | --- | --- |
| Path loss model |  | PLd [dB]=a⋅log10d+b+c+N |  |  |
| a |  | 21.1 |  | 46.9 |
| b |  | -10.5 |  | -80.0 |
| c |  | -0.9 |  | -0.9 |
| σN | 6.0 |  | 3.3 |  |

  - a and b : Coefficients of linear fitting

  - d : Tx-Rx distance in mm.

  - N : Normally distributed variable with standard deviation N

  - c:  The difference between the averaged signal levels of the dipole or colinear antenna used in the whole measurement and that of the chip antenna

# 8.2.5. Body surface to body surface CM3 (Scenario S4 & S5) for 900 MHz

8.2.5. A

The following path loss model is based on measurements that cover frequencies of 950-956 MHz. Further measurement set up, derivation and data analysis can also be found in [23]. The table below summarizes the model and corresponding parameters.

|  |  | Hospital Room |  | Anechoic Chamber |
| --- | --- | --- | --- | --- |
| Path loss model |  | PLd [dB]=a⋅log10d+b+c+N |  |  |
| a |  | 24.2 |  | 45.8 |
| b |  | -8.9 |  | -54.5 |
| c |  | -7.0 |  | -7.0 |
| σN | 3.9 |  | 8.3 |  |

  - a and b : Coefficients of linear fitting

  - d : Tx-Rx distance in mm.

  - N : Normally distributed variable with standard deviation N

  - c:  The difference between the averaged signal levels of the dipole or colinear antenna used in the whole measurement and that of the chip antenna

8.2.5. B

The following model is based on measurements at frequency of 915 MHz. Details of the measurement set up, derivation and data analysis can be found in [24]. The path loss follows an exponential decay around the perimeter of the body. It flattens out for large distance due to the contribution of multipath components from indoor environment. The table below summarizes the model and corresponding parameters.

| Path loss model |  | PL(d)[dB]= −10log10P0 e−m0d + P1 + σPnP |
| --- | --- | --- |
| P0 [dB] |  | -1.9 |
| M0 [dB/cm] |  | 2.1 |
| P1 [dB] |  | -59.4 |
| σp [dB] | 3.9 |  |

- P0 : The average loss close to the antenna

- M0 : The average decay rate in dB/cm for the surface wave traveling around the perimeter of the body

- P1 : The average attenuation of components in an indoor environment radiated away from the body and reflected back towards the receiving antenna

- p : The log-normal variance in dB around the mean, representing the variations measured at different body and room locations. This parameter will depend on variations in the body curvature, tissue properties and antenna radiation properties at different body locations.

The small scale fading is represented by a Ricean distribution with K factor that decreases as the path loss increases. The delay spread is normally distributed. The table below summarizes the model and corresponding parameters.

| Small-scale fading | KdB= K0 − mKPdB + σKnK |
| --- | --- |
| K0 [dB] | 40.1 |
| mk [dB] | 0.61 |
| σk [dB] | 2.4 |
| Parameters of the mean value of the delay spread |  |
| Distance [cm] | trms [ns] |
| 15 | 3 |
| 45 | 9 |
| Parameters of the 90% cumulative value of the delay spread |  |
| Distance [cm] | trms [ns] |
| 15 | 5 |
| 45 | 15 |

- K0 : The fit with measurement data for the K-factor for low path loss

- mk : The slope of the linear correlation between path loss and K-factor

- PdB : Path loss in dB

- k : The log-normal variance of the measured data between path loss and K-factor

- nk : A unit mean and variance Gaussian random variable ?

8.2.5. C

The following path loss measurements is at frequency of 820 MHz. Details of the measurement set up, derivation and data analysis can be found in [25]. The table below summarizes the results.

As observed, the dominant factor affecting fading in the channel appears to be the movement of the test subject. This is to be expected as movement causes the separation and orientation of the antennas to change.

| Path loss | PL[dB]= Ptx−Prx + Gamplifiers−Lcable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Receiver at right hip |  |  |  |  |  |  |  |  |  |  |  | Receiver at chest |  |  |  |  |  |
| Action | Chest |  | Right Wrist |  | Left Wrist |  | Right Ankle |  | Left Ankle |  | Back |  | Back |  | Right Wrist |  | Right Ankle |  |
| Standing | 57.4 |  | 50.2 |  | 59.8 |  | 54.3 |  | 68.7 |  | 61.8 |  | 66.3 |  | 54.5 |  | 54.3 |  |
| Walking |  | 52.9 |  | 38.4 |  | 63.6 |  | 48.1 |  | 55.5 |  | 57.1 |  | 63.8 |  | 51.3 |  | 56.9 |
| Running |  | 44.1 |  | 37.2 |  | 60.2 |  | 48.9 |  | 54.2 |  | 62.3 |  | 66.3 |  | 49.4 |  | 54.1 |

- Ptx: Transmitted power

- Prx: The RMS received power

- Gamplifiers: Amplifier gain

- Lcable: Cable loss

# 8.2.6. Body surface to body surface CM3 (Scenario S4 & S5) for 2.4 GHz

8.2.6. A

The following path loss model is based on measurements that cover frequencies of 2.4-2.5GHz. Details of the measurement set up, derivation and data analysis can be found in [23]. The table below summarizes the model and corresponding parameters.

|  |  | Hospital Room |  | Anechoic Chamber |
| --- | --- | --- | --- | --- |
| Path loss model |  | PLd [dB]=a⋅log10d+b+c+N |  |  |
| a |  | 8.32 |  | 46.4 |
| b |  | 37.2 |  | -49.4 |
| c |  | -7.5 |  | -7.5 |
| σN | 2.5 |  | 2.7 |  |

  - a and b : Coefficients of linear fitting

  - d : Tx-Rx distance in mm.

  - N : Normally distributed variable with standard deviation N

  - c:  The difference between the averaged signal levels of the dipole or colinear antenna used in the whole measurement and that of the chip antenna

8.2.6. B

The following model is based on measurements at frequency of 915 MHz. Details of the measurement set up, derivation and data analysis can be found in [24]. The path loss follows an exponential decay around the perimeter of the body. It flattens out for large distance due to the contribution of multipath components from indoor environment. The table below summarizes the model and corresponding parameters.

| Path loss model |  | PL(d)[dB]= −10log10P0 e−m0d + P1 + σPnP |
| --- | --- | --- |
| P0 [dB] |  | -25.8 |
| m0 [dB/cm] |  | 2.0 |
| P1 [dB] |  | -71.3 |
| σp [dB] | 3.6 |  |

- P0 : The average loss close to the antenna

- M0 : The average decay rate in dB/cm for the surface wave traveling around the perimeter of the body

- P1 : The average attenuation of components in an indoor environment radiated away from the body and reflected back towards the receiving antenna

- p : The log-normal variance in dB around the mean, representing the variations measured at different body and room locations. This parameter will depend on variations in the body curvature, tissue properties and antenna radiation properties at different body locations.

The small scale fading is represented by a Ricean distribution with K factor that decreases as the path loss increases. The delay spread is normally distributed. The table below summarizes the model and corresponding parameters.

| Small-scale fading |  | KdB= K0 − mKPdB + σKnK |
| --- | --- | --- |
| K0 [dB] | 30.6 |  |
| mk [dB] | 0.43 |  |
| σk [dB] | 3.4 |  |
| Parameters of the mean value of the delay spread |  |  |
| Distance [cm] | trms [ns] |  |
| 15 | 6 |  |
| 45 | 16 |  |
| Parameters of the 90% cumulative value of the delay spread |  |  |
| Distance [cm] | trms [ns] |  |
| 15 | 11 |  |
| 45 | 22 |  |

- K0 : The fit with measurement data for the K-factor for low path loss

- mk : The slope of the linear correlation between path loss and K-factor

- PdB : Path loss in dB

- k : The log-normal variance of the measured data between path loss and K-factor

- nk : A unit mean and variance Gaussian random variable ?

8.2.6. C

The following path loss measurements is at frequency of 2.36 GHz. Details of the measurement set up, derivation and data analysis can be found in [25]. The table below summarizes the results.

As observed, the dominant factor affecting fading in the channel appears to be the movement of the test subject. This is to be expected as movement causes the separation and orientation of the antennas to change.

| Path loss | PL[dB]= Ptx−Prx + Gamplifiers−Lcable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Receiver at right hip |  |  |  |  |  |  |  |  |  |  |  | Receiver at chest |  |  |  |  |  |
| Action | Chest |  | Right Wrist |  | Left Wrist |  | Right Ankle |  | Left Ankle |  | Back |  | Back |  | Right Wrist |  | Right Ankle |  |
| Standing | 65.3 |  | 44.5 |  | 74.7 |  | 60.9 |  | 70.7 |  | 75.3 |  | 73 |  | 70.5 |  | 66.3 |  |
| Walking |  | 59.1 |  | 47.3 |  | 59.8 |  | 53.9 |  | 58.5 |  | 67.4 |  | 72 |  | 64.9 |  | 62.4 |
| Running |  | 55.9 |  | 36.3 |  | 52.5 |  | 55 |  | 59 |  | 68.5 |  | 71.7 |  | 57.4 |  | 63.3 |

- Ptx: Transmitted power

- Prx: The RMS received power

- Gamplifiers: Amplifier gain

- Lcable: Cable loss

# 8.2.7. Body surface to body surface CM3 (Scenario S4 & S5) for 3.1-10.6 GHz

8.2.7. A

The following path loss model is based on measurements that cover frequencies of 3.1-10.6 GHz. Measurement set up, derivation and data analysis can be found in [23]. The table below summarizes the model and corresponding parameters.

|  |  | Hospital Room |  | Anechoic Chamber |
| --- | --- | --- | --- | --- |
| Path loss model |  | PLd [dB]=a⋅log10d+b+N |  |  |
| a |  | 8.43 |  | 17.0 |
| b |  | 31.8 |  | 9.8 |
| σN | 2.8 |  | 4.66 |  |

  - a and b : coefficients of linear fitting

  - d : Tx-Rx distance in mm

  - N : Normally distributed variable with zero mean and standard deviation N

8.2.7. B

The following path loss model is based on measurements that cover frequencies of 3.1-10.6 GHz. Measurement set up, derivation and data analysis can be found in [24]. The table below summarizes the corresponding parameters.

| Around torso |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Antenna separation from body surface |  | 0 mm |  |  | 5 mm |  |  | 10 mm |
| P0 [dB] |  | 56.1 |  |  | 48.4 |  |  | 45.8 |
| d0 [m] |  | 0.1 |  |  | 0.1 |  |  | 0.1 |
| n | 5.8 |  |  | 5.9 |  |  | 6.0 |  |
| Along torso |  |  |  |  |  |  |  |  |
| Antenna separation from body surface |  |  | 0 mm |  |  | 5 mm |  |  |
| P0 [dB] |  |  | 56.5 |  |  | 44.6 |  |  |
| d0 [m] |  |  | 0.1 |  |  | 0.1 |  |  |
| n |  |  | 3.1 |  |  | 3.1 |  |  |

- n: Path loss exponent

- P0: TBD

- d0: TBD

# 8.2.8. Body surface to external CM4 (Scenario S6 & S7) for 900 MHz

The following path loss measurements is at frequency of 820 MHz. Details of the measurement set up, derivation and data analysis can be found in [26]. The table below summarizes the results.

The transmitter is placed on the chest and the receiver is away from the body with various distances as outlined.

|  | LOS |  |  |  |  |  | NLOS |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Distance (m) | 1 |  | 2 |  | 3 | 4 | 1 | 2 | 3 | 4 |
| Standing | 5.187 |  | 14.037 |  | 16.611 | 11.492 | 20.626 | 18.356 | 26.718 | 20.122 |
| Walking | 4.481 | 9.567 |  | 16.395 |  | 16.755 | 7.437 | 19.416 | 25.272 | 25.646 |

# 8.2.9. Body surface to external CM4 (Scenario S6 & S7) for 2.4 GHz

The following path loss measurements is at frequency of 2.36 GHz. Details of the measurement set up, derivation and data analysis can be found in [26]. The table below summarizes the results.

The transmitter is placed on the chest and the receiver is away from the body with various distances as outlined.

|  | LOS |  |  |  |  |  | NLOS |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Distance (m) | 1 |  | 2 |  | 3 | 4 | 1 | 2 | 3 | 4 |
| Standing | 17.068 |  | 16.070 |  | 18.602 | 27.576 | 23.989 | 31.562 | 22.851 | 26.184 |
| Walking | 7.991 | 14.269 |  | 15.243 |  | 20.930 | 22.137 | 33.266 | 29.863 | 31.208 |

# 8.2.10. Body surface to external CM4 (Scenario S6 & S7) for 3.1-10.6 GHz

TBA

# 9. References

- ERC Recommendation 70-03 relating to the use of Short Range Device (SRD), European Conference of Postal and Telecommunications Administrations, CEPT/ERC 70-03, Tromsø, Norway, 1997.

- FCC, Medical implant communications, January 2003,

http://wireless.fcc.gov/services/index.htm?job=service_home&id=medical_implant

- W.-T. Chen; H.-R. Chuang, “Numerical computation of human interaction with arbitrarily oriented superquadric loop antennas in personal communications,” IEEE Trans. on Antenna and Propagation, vol.46, no. 6, pp. 821-828, June 1998.

- Kamya Y. Yazdandoost and Ryuji Kohno, “The Effect of Human Body on UWB BAN Antennas,” IEEE802.15-07-0546-00-0ban.

- Kamya Y. Yazdandoost and Ryuji Kohno, “Wireless Communications for Body Implanted Medical Device,” Asia Pacific Microwave Conference, APMC2007, pp.

- Kamya Y. Yazdandoost et al, “Channel Characterization for BAN Communications,” IEEE802.15-07-0641-00-0ban.

- Andreas F. Molisch et al, “A Comprehensive Model for Ultrawideband Propagation Channels,” IEEE Global Telecommunications Conference, GLOBECOM '05. Vol.6, pp. 3648-3653.

- 15-07-0939-01-0ban-ieee-802-15-6-regulation-subcommittee-report

- 15-07-0735-06-0ban-ban-application-matrix_amaledit

- C. H. Duney, H. Massoudi, and M. F. Iskander, “Radiofrequency radiation dosimetry handbook,” USAF School of Aerospace Medicine, October 1986.

- C. Gabriel and S. Gabriel, “Compilation of the dielectric properties of body tissues at RF and microwave frequencies,” AL/OE-TR-1996-0037, June 1996, http://www.brooks.af.mil/AFRL/HED/hedr/reports/dielectricReport/Report.html.

- Italian National Research Council, Institute for Applied Physics, “Dielectric properties of body tissues,” http://niremf.ifac.cnr.it

- P. Gandhi, “.Biological effects and medical applications of electromagnetic energy,” Prentice Hall, Englewood Cliffs, N.J., 1990.

- E. Reusens, W. Joseph, G. Vermeeren, and L. Martens, „On-body measurements and characterization of wireless communication channel fro arm and torso of human,“ International Workshop on Wearable and Implantabel Body Sensor Networks, BSN07, Achen, March 2007, pp. 26-28.

- A. Fort, J. Ryckaert, C. Desset, P. De Doncker, P. Wambacq, and L. Van Biesen, “ Ultra-wideband channel model for communication around the human body, ” IEEE Journal on Selected Areas in Communications, vol. 24, pp.927-933, April 2006.

- Bernard Sklar, “Rayleigh fading channels in mobile digital communication system part I: characterization”, IEEE Communication Magazine, pp.90-100, July 1997.

- Kamya Yekeh Yazdandoost and Ryuji Kohno, “Antenna for Medical Implanted Communications System,” IEEE 802.15-07-0785-00-0ban, July 2007.

- TBD

- John Hagedorn, Judith Terrill, Wenbin Yang, Kamran Sayrafian, Kamya Yazdandoost, Ryuji Kohno, “MICS Channel Characteristics; Preliminary Results”, IEEE 802.15-08-0351-00-0006, September 2008.

- John Hagedorn, Judith Terrill, Wenbin Yang, Kamran Sayrafian, Kamya Yazdandoost, Ryuji Kohno, “A Statistical Path Loss Model for MICS”, IEEE 802.15-08-0519-01-0006, September 2008.

- Arthur Astrin, “Measurements of body channel at 13.5 MHz,” IEEE 802.15-08-0590-00-0006, August 2008.

- Jung-Hwan Hwang, Il-Hyoung Park, and Sung-Weon Kang, “Channel model for human body communication,” IEEE 802.15-08-0577-00-0006, August 2008.

- Takahiro Aoyagi, Jun-ichi Takada, Kenichi Takizawa, Norihiko Katayama, Takehiko Kobayashi, Kamya Yekeh Yazdandoost, Huan-bang Li and Ryuji Kohno, “Channel model for wearable and implantable WBANs,” IEEE 802.15-08-0416-01-0006, July 2008.

- Guido Dolmans and Andrew Fort, “Channel models WBAN-Holst centre/IMEC-NL,” IEEE 802.15-08-0418-01-0006, July 2008.

- Dino Miniutti, Leif Hanlen, David Smith, Andrew Zhang, Daniel Lewis, David Rodda, Ben Gilbert, “Narrowband channel characterization for body area network,” IEEE 802.15-08-0421-00-0006, July 2008.

- Dino Miniutti, Leif Hanlen, David Smith, Andrew Zhang, Daniel Lewis, David Rodda, Ben Gilbert, “Narrowband on body to off body channel characterization for ban,” IEEE 802.15-08-0559-00-0006, August 2008.