IEEE 802.15

Wireless Personal Area Networks

| Project | IEEE 802.15 Working Group for Wireless Personal Area Networks |  |
| --- | --- | --- |
| Title | IEEE 802.15.6 Regulation Subcommittee Report |  |
| Date Submitted | May 2010 |  |
| Source | Huan-Bang Li Jean Schwoerer YangMoon Yoon John Farserotu Wenbin Yang Kamran Sayrafian Dino Miniutti Daniel Lewis Andrew Gowans Jeremy Moss | [lee@nict.go.jp] [jean.schwoerer@orange-ftgroup.com] [ymyoon@korpa.or.kr] [john.farserotu@csem.ch] [wyang@nist.gov] [ksayrafian@nist.gov] [dino.miniutti@nicta.com.au] [Daniel.Lewis@nicta.com.au] [Andrew.Gowans@ofcom.org.uk] [j.moss@philips.com] |
| Re: |  |  |
| Abstract | This document summarizes the review and discussion results for regulation issues for medical body area networks. |  |
| Purpose | [This document is intent to help TG6’s discussion.] |  |
| Notice | This document has been prepared to assist the IEEE 802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |  |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by 802.15. |  |

TABLE OF CONTENTS

1.	Introduction	5

2.	References	6

3.	AUSTRALIA	7

3.1.	Low Interference Potential Devices	7

3.1.1.	Low Interference Potential Devices (LIPD )	7

3.1.2.	Conditions of Operation	7

3.1.2.1.	Managing Interference	7

3.1.2.2.	Compliance with Standards	8

3.1.2.3.	Operating Frequencies	8

3.2.	UWB Frequency Band	12

4.	China	14

4.1.	UWB Frequency Band	14

5.	Europe	16

5.1.	Medical and General SRD Frequency Bands	17

5.1.1.	Inside CEPT	17

5.1.1.1.	Medical devices	17

5.1.1.2.	General SRD	18

5.1.2.	Regulations in the European Union	21

5.1.3.	Relevant ETSI Harmonised Standards	21

5.1.3.1.	Medical Frequency Bands	21

5.1.3.2.	General Non- Specific SRD’s	22

5.2.	UWB Frequency Band	23

5.2.1.	Inside CEPT	23

5.2.2.	Inside the European Union (27 countries who are also CEPT Member)	26

5.2.3.	Relevant ETSI Harmonised Standards	29

6.	Japan	30

6.1.	Medical Frequency Bands	30

6.1.1.	Wireless Medical Telemetry System (WMTS)	30

6.1.2.	Medical Implant Communication Systems (MICS)	30

6.2.	UWB Frequency Band	32

6.2.1.	Spectrum Mask	32

6.2.2.	General Technical Requirements	32

6.2.2.1.	General Conditions	32

6.2.2.2.	Transmitter	33

6.2.2.3.	Receiver	36

6.2.2.4.	Controller	39

6.2.2.5.	Connection with Telecommunication Circuit Equipment	39

6.2.2.6.	Antenna	40

6.2.2.7.	Others	40

7.	Korea	42

7.1.	Medical Frequency Bands	42

7.2.	UWB Frequency Band	44

8.	UK	46

8.1.	General Non-Specific Short Range Devices	46

8.2.	Industrial / Commercial Telemetry and Telecommand	48

8.3.	Medical and Biological Applications	49

9.	USA	51

9.1.	Medical Frequency Bands	51

9.1.1.	Wireless Medical Telemetry System (WMTS)	51

9.1.1.1.	WMTS Transmitter frequencies	51

9.1.1.2.	General Technical Requirements	51

9.1.1.3.	Radiation Testing	53

9.1.2.	Medical Device Radiocommunications Service	53

9.1.2.1.	Summary of MedRadio operations rules	55

9.1.2.2.	Equipment and Interference Safeguards	55

9.1.2.3.	Radiation Testing	56

9.2.	UWB Frequency Band	57

9.2.1.	Definitions	57

9.2.2.	General technical requirements	57

10.	ISM Band	60

11.	Related regulation for HBC	62

11.1.	EU	62

11.2.	Japan	62

11.3.	Korea	63

11.4.	UK	64

11.5.	USA	65

11.5.1.	FCC Regulations for low-power, non-licensed transmitters	65

11.5.2.	FCC Regulations for low-power, non-licensed transmitters	65

12.	Summary	67

# Introduction

This is the final report of BAN Regulatory Subcommittee of the TG15.6 for body area networks (BAN). This document provides a summary of frequency regulations for medical applications in different countries and regions based on a number of review works presented (see References).

The document is divided into ten sections. Some related references are given in section 2. Regulations in Australia, China, EU, Korea, Japan, UK, USA, as well as the ISM band regulation, are described respectively from sections 3 to 10. The document is shortly summarized in section 11.

The intention of this document is to help TG15.6 members on proposing and selecting the appropriate PHY.

# References

| Ref. | [date(yyyy/mm)] | DCN | Title |
| --- | --- | --- | --- |
| [1] | [2006/07] | 15-06-0340-00-0ban |  |
| [2] | [2007/03] | 15-07-0640-01-0ban |  |
| [3] | [2007/05] | 15-07-0722-00-0ban |  |
| [4] | [2007/05] | 15-07-0728-00-0ban |  |
| [5] | [2007/09] | 15-07-0871-00-0ban |  |
| [6] | [2007/11] | 47 CFR 95.601-95.673 Subpart E |  |
| [7] | [2007/11] | 47 CFR 95.628(a) |  |
| [8] | [2008/07] | 47 CFR Part 15 Section 15.209 |  |
| [9] | [2010/05] | 15-10-0318-00-0006 |  |

# AUSTRALIA

This section describes the relevant information on the Australian Regulations pertaining to BAN usage.

## Low Interference Potential Devices

Radiocommunications devices authorised under class licences are typically low power transmitters providing short-range communications that do not require individual frequency coordination for interference management purposes.

### Low Interference Potential Devices (LIPD )

The LIPD Class Licence authorises users to operate a wide range of low power radiocommunications devices in various segments of the radiofrequency spectrum. The class licence sets out the conditions under which many types of short-range devices may operate. These conditions always cover frequency bands of operation and radiated power limits. Other conditions are applied as necessary.

These lower power transmitters for short range communications do not require individual frequency coordination for interference management purposes. Examples of equipment covered by the LIPD Class Licence include garage door openers, home detention monitoring equipment, spread spectrum devices and personal alarms.

### Conditions of Operation

#### Managing Interference

A device operating under the LIPD Class Licence does so on an uncoordinated basis. It operates in spectrum that is used by other devices; that is, it shares the spectrum with other devices. The potential for interference between devices is managed by placing appropriate limits on the operating parameters of the devices (such as the type of device, radiated power levels, and areas and frequencies of operation).

Although in some circumstances LIPDs can be used for radio applications with commercial or safety-of-life implications, users of such applications are encouraged to pay particular regard to the suitability of operating under this class licence for their radiocommunications needs.

Should interference occur, the onus is on the user of a LIPD to take measures to resolve that interference, for example by retuning or ceasing to operate the device. Some LIPDs are designed to have the capacity to be retuned in order to assist the user in avoiding local interference.

LIPDs operating in bands designated for industrial, scientific and medical (ISM) applications will not be afforded protection from interference which may be caused by ISM applications (e.g. microwave ovens). The ISM bands are:

- 13553-13567 kHz

- 26957-27283 kHz

- 40.66-40.70 MHz

- 918-926 MHz

- 2400-2500 MHz

- 5725-5875 MHz

- 24-24.25 GHz

#### Compliance with Standards

Over recent years concerns have been raised about the long-term health consequences of the use of radiocommunications transmitters. The Radiocommunications (Electromagnetic Radiation - Human Exposure) Standard 2001, which limits human exposure to electromagnetic radiation from radiocommunications transmitters, came into effect on 1 January 1999. From 22 November 2000, cellular mobile telecommunications handsets and cordless telephones and cradles that are capable of operating in the frequency range 800 MHz to 2.5 GHz, and operate under the LIPD Class Licence, must comply with that standard.

#### Operating Frequencies

This is a trimmed down list of “general purpose” frequencies that are most suitable for BAN usage. It also includes frequency bands suitable for medical uses; they have been shaded with a grey background.

| Item | Class of transmitter | Permitted operating frequency band (MHz) (lower limit exclusive, upper limit inclusive) | Maximum EIRP | Limitations |
| --- | --- | --- | --- | --- |
| 1 | All transmitters | 0.000–0.014 | 200 µW |  |
| 2 | All transmitters | 0.014–0.01995 | 50 µW |  |
| 3 | All transmitters | 0.02005–0.07 | 7.5 µW |  |
| 4 | All transmitters | 0.07–0.16 | 3 µW |  |
| 5 | All transmitters | 1. 0.16–0.285 2. 0.325–0.415 | 500 nW |  |
| 6 | All transmitters | 3.025–3.155 | 7.5 nW |  |
| 7 | All transmitters | 3.5–3.7 | 30 pW |  |
| 8 | All transmitters | 1. 3.7–3.95 2. 4.438–4.65 | 7.5 nW |  |
| 9 | All transmitters | 13.553–13.567 | 100 mW |  |
| 10 | All transmitters | 24–24.89 | 10 mW |  |
| 11 | All transmitters | 26.957–27.283 | 1 W | 1. Separation of the operating frequency from the centre frequency of any adjacent citizen band radio channel must be at least 5 kHz. 2. The emission bandwidth must not exceed 10 kHz. |
| 12 | All transmitters | 1. 29.7–29.72 2. 30–30.0625 3. 30.3125–31 4. 36.6–37 5. 39–39.7625 6. 40.25–40.66 | 100 mW |  |
| 13 | All transmitters | 40.66–41 | 1 W |  |
| 14 | All transmitters | 54–56 | 2.5 mW |  |
| 15 | All transmitters | 1. 70–70.24375 2. 77.29375–77.49375 3. 150.7875–152.49375 4. 173.29375–174 | 100 mW |  |
| 16 | All transmitters | 1. 225–242 2. 244–267 3. 273–303.95 4. 304.05–328.6 5. 335.4–399.9 | 10 µW |  |
| 17 | All transmitters | 433.05–434.79 | 25 mW |  |
| 18 | All transmitters | 915–928 | 3 mW |  |
| 19 | All transmitters | 2400–2483.5 | 10 mW |  |
| 20 | All transmitters | 1. 10500–10550 2. 24000–24250 | 100 mW |  |
|  |  |  |  | 4. When transmitting in an unused broadcasting services bands channel, and in the coverage area of a broadcasting station (including a repeater or translator station) operating in an adjacent channel, the channel centre frequency of the wireless audio transmitter must be at least 400 kHz above the upper edge of the adjacent channel, or 400 kHz below the lower edge of the adjacent channel. |
| 23 | Biomedical telemetry transmitters | 174–230 | 10 µW |  |
| 24 | Biomedical telemetry transmitters | 520–668 | 3 mW | Transmission in a TV channel must not originate in the licence area of an analogue TV broadcasting station (including a repeater or translator station) operating in the same channel. |
| 45 | Digital modulation transmitters | 915–928 | 1 W | 1. The radiated peak power spectral density in any 3 kHz is limited to 25 mW per 3 kHz. 2. The minimum 6 dB bandwidth must be at least 500 kHz. |
| 45A | Digital modulation transmitters | 2400–2483.5 | 4 W | 1. The radiated peak power spectral density in any 3 kHz is limited to 25 mW per 3 kHz. 2. The minimum 6 dB bandwidth must be at least 500 kHz. |
| 45B | Digital modulation transmitters | 5725–5850 | 4 W | 1. The radiated peak power spectral density in any 3 kHz is limited to 25 mW per 3 kHz. 2. The minimum 6 dB bandwidth must be at least 500 kHz. |
| 49 | Medical implant communications systems transmitters | 402–405 | 25 µW (averaged over the transmission burst within a reference bandwidth of 300 kHz) | 1. The maximum EIRP applies outside the body. 2. Systems must have a minimum of nine channels selectable by the system controller and spread across the whole band. 3. Implanted transmitters must only transmit under external control, except for medical implant events. 4. Systems must utilise a listen-before-transmit protocol. |
|  | Note 1 The systems and associated medical implant communications systems transmitters mentioned in item 48 are devices that require marketing approval by the Therapeutic Goods Administration. Note 2 A medical implant event is an occurrence or lack of occurrence, recognised by a medical implant device or a health care professional, that requires the immediate transmission of data by the medical implant communications systems transmitter to protect the safety or wellbeing of the person that the medical implant device has been implanted. |  |  |  |
| 50 | Medical implant telemetry systems transmitters | 403.560–403.760 | 100 nW | The maximum EIRP applies outside the body. |
|  | Note The systems and associated medical implant devices mentioned in item 49 are devices that require marketing approval by the Therapeutic Goods Administration. |  |  |  |
| 52 | Frequency hopping transmitters | 915–928 | 1 W | A minimum of 20 hopping frequencies must be used. |
| 53 | Frequency hopping transmitters | 2400–2483.5 | 500 mW | A minimum of 15 hopping frequencies must be used. |
| 54 | Frequency hopping transmitters | 2400–2483.5 | 4 W | A minimum of 75 hopping frequencies must be used. |
| 55 | Frequency hopping transmitters | 5725–5850 | 4 W | A minimum of 75 hopping frequencies must be used. |

## UWB Frequency Band

Ultra Wideband (UWB) technology generally involves the radiation, reception and processing of very wide bandwidth radiofrequency emissions for short-range applications.  UWB applications include automotive collision-avoidance systems, and high data rate interference-tolerant communications.

Typically, the emissions from a UWB transmitter will span a number of radiofrequency bands that have been allocated for a range of different purposes.  For example, a 24 GHz UWB transmitter might occupy a bandwidth around 5 GHz, which would span frequencies used for such purposes as microwave fixed links, space research, radio astronomy, amateur radio and satellite communications - for details on current spectrum allocations see the Australian Radiofrequency Spectrum Plan.

As a general principle, ACMA seeks where possible to align Australia's spectrum management arrangements with those of the rest of the world, for the trade and community benefits that brings. In this context, by early 2003 the development of an international regulatory framework generally applicable to UWB was at an early stage. Two general concerns were raised internationally to be resolved:

- the potential for interference to the services already using those bands; and

- the regulatory method for accommodating these devices within the radio regulations (UWB is a group of technologies, not a service as defined by the ITU).

Carriage of these issues within the framework of the International Telecommunication Union (ITU) has been given to a specific task group (TG 1/8) formed within Study Group 1 in 2002. The work of the task group concluded in October 2005 and several recommendations have now been finalised through the ITU-R approval process.

Additionally, ACMA and DCITA have been working on various licensing and legislative issues that will need to be resolved before, generally speaking, ACMA will have sufficiently flexible licensing powers to be able to authorise the use of UWB devices in this country.  The ACMA paper Ultra Wideband (UWB) - A Background Brief (PDF 252 kb), released in May 2003, provides useful information on these matters.

In 2005 arrangements were introduced to allow the use of UWB technology in scientific trials in Australia:

- Media Release: Use of ultra wideband approved for the first time

Studies were completed in early 2006 into providing support for the introduction of 24 GHz Ultra-Wideband Short-Range Radars for Automotive Applications (PDF 722 kb).

In July 2006 arrangements were introduced to allow the use of 24 GHz Ultra-Wideband short range radar devices for automotive use in Australia:

- Media Release: Licensing for anti-collision vehicle radar.

# China

This section describes the frequency regulation on UWB in China.

## UWB Frequency Band

UWB regulation on UWB was issued by Ministry of Industry and Information Technology, China in December 6, 2008. The Main Contents for the Regulation on the UWB are summarized as follows.

a.  Transmission power at -10dB bandwidth must be equal to or larger than 500MHz.

b.  UWB emission limits in e.i.r.p. for UWB are shown in Table 4.1

Table 4.1 UWB emission limits in e.i.r.p.

| Frequency (GHz) | e.i.r.p.（dBm/MHz） | Detection method |
| --- | --- | --- |
| < 1.6 | -90 | RMS |
| 1.6－3.6 | -85 |  |
| 3.6 – 6.0note 1 | -70 |  |
| 6.0 – 9.0 | -41 |  |
| 9.0－10.6 | -70 |  |
| >10.6 | -85 |  |

Note 1 :  Up to the date 31 Dec. 2010 , the UWB emission limit for 4.2 - 4.8GHz could be -41.0dBm/MHz. After the date above mentioned, the UWB equipment shall have detection and interference avoided technology for which the technology availability should be approved by the China Radio Regulatory Organization .

c. Spurious Emission limits are shown in Table 4.2

Table 4.2 Spurious emission limit

| Transmitter status | 48.5 - 72.5 MHz 76 - 108 MHz 167 - 223 MHz 470 - 798 MHz (measuring BW=100kHz） | Other frequency bands on 30 MHz - 1GHz (measuring BW=100kHz) | 1 - 40 GHz (measuring BW=1MHz) |
| --- | --- | --- | --- |
| Transmit | -54 dBm | -36 dBm | -30 dBm |
| Stanby | -57 dBm | -57 dBm | -47 dBm |

d.  It is prohibited to use UWB Equipment in Aircraft

e. According to the Note “ CHN 12” listed  in the “ Regulations on Radio Frequency Allocation , People’s Republic of China , Dec. 2006 )， it is prohibited to use UWB Equipment in around 1km of the radio-astronomy station

f. The management for UWB Equipment is following the Micro-power (Short range) Radio Tx. Equipment Radio Regulatory Management Requirements in China . UWB equipment shall obtain the type approval certification for the Tx equipment by MIIT before putting into service.

g. UWB Tx equipment shall not create any interference to other radio service stations. UWB shall not ask any interference protection other radio service stations .

# Europe

This section describes the frequency regulation for medical applications in Europe. It should be remembered that as well as the individual national spectrum regulators in Europe there are also three different organizations who produce regulatory deliverables for Spectrum. These deliverables all have different legal standings in their applicability. The EU, CEPT and ETSI.

- Conference of European Post & Telecommunication (CEPT)

The CEPT is a voluntary organization that consists of 48 members. All of the 27 member states of the EU are members of CEPT. Other members include CEPT has 3 main deliverables ECC(formerly ERC) Decisions, Recommendations and Reports. ECC Decisions are regulatory texts providing measures on significant harmonisation matters, which CEPT member NRAs are strongly urged to follow. As any other CEPT deliverables, ECC Decisions are not obligatory legislative documents; however, they are normally implemented by many CEPT administrations. To see the implementation status go to http://www.erodocdb.dk/doks/doccategoryECC.aspx?doccatid=1 and click on ‘implementation overview’. If there is no implementation icon, this means that the Decision has not been implemented in any administration.

- European Union (EU)

In the European Union the deliverables are produced either as part of an EU Directive or as an EC Spectrum Decision. Both of these deliverables are mandatory for the 27 member states of the European Union.

- European Telecom Standards Institute (ETSI)

In order to place products on the EU Marketplace they must comply with the minimum requirements of the article 3.2 of R&TTE directive. ETSI produce Harmonised Standards (HS or EN) which is a regulatory deliverable which is published in the Official Journal of the EU. Declaring compliance with a HS is the preferred option for the majority of vendors when placing products on the EU market. The ETSI HS is also accepted by other CEPT countries as part of their equipment compliance procedures.

## Medical and General SRD Frequency Bands

### Inside CEPT

#### Medical devices

(ERC/DEC/(01)17) was the first decision concerning Ultra Low Power Medical Implant. It as been decided :

- to designate the frequency band 402 - 405 MHz for the use of equipment for Medical Implants which comply with the technical characteristics shown in Table 5.1 below;

- to exempt Short Range Devices for Medical Implants covered by this Decision from individual licensing;

Table 5.1 Technical characteristics for implant device

It should be noted that :

- Those bands are reserved only for medical and implanted system

- Therefore they can't be used for on-body communication (even for medical applications)

- SRDs in general operate in shared bands and are not permitted to cause harmful interference to other radio services; that in general SRDs cannot claim protection from other radio services

- The 401-402 MHz and the 405-406 MHz frequency bands were selected for the emerging medical implant technologies, using Ultra Low Power – Active Medical Implants (ULP-AMI) and ULP-AMI-P (peripheral devices for ULP-AMI).

Those bands allocation and the existing shortwave band allocation can be summarized in the following Table 5.2.

Table 5.2 Medical bands in 70-03

Note 1: Due to the limited available spectrum of 1 MHz, a maximum bandwidth of 100 kHz is proposed for these bands to ensure that several users could access the band concurrently.

Note 2: Systems not providing frequency agility based on ambient RF field sensing, be limited to a maximum permitted e.r.p. of 250 nano watts with a duty cycle of ≤0.1%.

#### General SRD

ERC/REC 70-03

The Table 5.3 below is from annex 1 of ERC/REC 70-03 and it shows the bands that have been allocated for non specific (or generic) SRD use that could also be used for Body Area Networks (BAN) medical or otherwise.

Table 5.3 General SRD allocations from 70-03

|  | Frequency | Band Power / Magnetic Field | Duty Cycle | Channel spacing | ECC/ERC Decision | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| a | 6765-6795 kHz | 42 dBμA/m at 10m | No Restriction | No spacing |  |  |
| b | 13.553-13.567 MHz | 42 dBμA/m at 10m | No Restriction | No spacing |  |  |
| c | 26.957-27.283 MHz | 42 dBμA/m at 10m 10mW e.i.r.p. | No Restriction | No spacing | ERC/DEC/(01)02 |  |
| d | 40.660-40.700 MHz | 10 mW e.r.p. | No Restriction | No spacing | ERC/DEC/(01)03 |  |
| e | 138.20-138.45 MHz | 10 mW e.r.p. | < 1.0 % | No spacing |  |  |
| f | 433.050-434.790 MHz ( see note 4) | 10 mW e.r.p. | < 10 % | No spacing | ECC/DEC/(04)02 |  |
| f1 | 433.050-434.790 MHz (note 4bis) | 1 mW e.r.p. -13 dBm/10 kHz | up to 100% | No spacing | ECC/DEC/(04)02 | Power density limited to -13 dBm/10 kHz for wideband modulation with a bandwidth greater than 250 kHz |
| f2 | 434.040-434.790 MHz (note 4bis) | 10 mW e.r.p. | up to 100% | Up to 25 kHz | ECC/DEC/(04)02 |  |
| g | 863-870 MHz (note 3, 4 and 6) | ≤ 25 mW e.r.p. | ≤ 0.1% or LBT (note 1 and 5) | ≤ 100 kHz for 47 or more channels (note 2) |  | FHSS modulation |
|  |  | ≤ 25 mW e.r.p. (note 6) Power density : - 4.5 dBm/100 kHz (note 8) | ≤ 0.1% or LBT (note 1, 5 and 6) | No spacing |  | DSSS and other wideband modulation other than FHSS |
|  |  | ≤ 25 mW e.r.p. | ≤ 0.1% or LBT (note 1 and 5) | ≤ 100 kHz, for 1 or more channels (note 2 and 7) |  | Narrow /wide-band modulation |
| g1 | 868.000-868.600 MHz (note 4) | ≤ 25 mW e.r.p. | ≤ 1% or LBT (note 1) | No spacing, for 1 or more channels (note 2) | ERC/DEC/(01)04 | Narrow/wide-band modulation No channel spacing, however the whole stated frequency band may be used |
| g2 | 868.700-869.200 MHz (note 4) | ≤ 25 mW e.r.p. | ≤ 0.1% or LBT (note 1) | No spacing, for 1 or more channels (note 2) | ERC/DEC/(01)04 | Narrow/ wide-band modulation No channel spacing, however the whole stated frequency band may be used |
| g3 | 869.400-869.650 MHz (note 4) | ≤ 500 mW e.r.p. | ≤ 10% or LBT. (note 1) | 25 kHz (for 1 or more channels) | ERC/DEC/(01)04 | Narrow/ wide-band modulation The whole stated frequency band may be used as 1 channel for high speed data transmission |
| g4 | 869.700-870.000 MHz (note 4bis) | ≤ 5 mW e.r.p. | up to 100% | No spacing (for 1 or more channels) | ERC/DEC/(01)04 | Narrow /wide-band modulation No channel spacing, however the whole stated frequency band may be used |
| h | 2400.0-2483.5 MHz | 10 mW e.i.r.p. | No Restriction | No spacing |  |  |
| i | 5725-5875 MHz | 25 mW e.i.r.p. | No Restriction | No spacing | ERC/DEC/(01)06 |  |
| j | 24.00-24.25 GHz | 100 mW e.i.r.p. | No Restriction | No spacing |  |  |
| k | 61.0-61.5 GHz | 100 mW e.i.r.p. | No Restriction | No spacing |  |  |
| l | l 122-123 GHz | 100 mW e.i.r.p. | No Restriction | No spacing |  |  |
| m | m 244-246 GHz | 100 mW e.i.r.p. | No Restriction | No spacing |  |  |

Note 1: For frequency agile devices the duty cycle limit applies to the total transmission unless LBT is used. For LBT devices without frequency agility, the duty cycle limit applies.

Note 2: The preferred channel spacing is 100 kHz allowing for a subdivision into 50 kHz or 25 kHz.

Note 3: Sub-bands for alarms are excluded (see ERC/REC 70-03 Annex 7).

Note 4: The duty cycle, LBT or equivalent technique shall not be user dependent and shall therefore be guaranteed by appropriate technical means.

Note 4bis: Audio applications should be excluded. Voice applications allowed with spectrum access technique such as LBT or equivalent technique, the transmitter shall include a power output sensor controlling the transmitter to a maximum transmit period of 1 minute.

Note 5: Duty cycle may be increased to 1% if the band is limited to 865-868 MHz.

Note 6: For other wide-band modulation than FHSS and DSSS with a bandwidth of 200 kHz to 3 MHz, duty cycle can be increased to 1% if the band is limited to 865-868 MHz and power to ≤10 mW e.r.p.

Note 7: For other narrow-band modulation with a bandwidth of 50 kHz to 200 kHz, the band is limited to 865.5-867.5 MHz.

Note 8: The power density can be increased to +6.2 dBm/100 kHz and +0.8 dBm/100 kHz, if the band of operation is limited to 865-868 MHz and 865-870 MHz respectively.

Further notes on Frequency issues related to Table 2 above.

The bands in table 2  a - b - c - d f - f1 - f2 - h - i - j - k - l and m are also designated for industrial, scientific and medical (ISM) applications as defined in ITU Radio Regulations.For subband g) in table 2

Certain channels may be occupied by RFID operating at higher powers (See Annex 11 of ERC Rec. 70-03 for further details). To minimise the risk of interference from RFID, SRDs should use LBT with AFA or observe suitable separation distances. (In the high power RFID channels typically these may vary from 918 m (indoor) to 3.6 km (rural outdoor). In the remaining 2.2 MHz, where tags at -20 dBm e.r.p. occupy the spectrum, this may vary from 24 m (indoor) to 58 m (rural outdoor)). The adjacent frequency band above 870 MHz has been designated for use by the high powered TETRA and other digital land mobile PMR/PAMR systems. Manufacturers should take this into account in the design of equipment and choice of power level Inside EU.

### Regulations in the European Union

The relevant regulations for Short Range Devices in the European Union are contained in the EC Decision 2006/771/EC on harmonisation of the radio spectrum for use of Short Range Devices. This decision also contains any frequencies allocated for medical SRD use.

For further details see http://www.erodocdb.dk/Docs/doc98/official/pdf/2006771EC.PDF . This decision has been updated recently by EC Decision 2008/432/EC see the following web address for full details http://www.erodocdb.dk/Docs/doc98/official/pdf/2008432EC.PDF .

### Relevant ETSI Harmonised Standards

#### Medical Frequency Bands

The designations used in Table 5.4 below (a to e) are cross reference with those used in table 1 above.

Table 5.4 ETSI standards for medical devices

|  | Frequency Band | ETSI Standard |
| --- | --- | --- |
| a | 402 - 405 MHz | EN 301 839 |
| a 1 | 401 - 402 MHz | EN 302 537 |
| a 2 | 405 - 406 MHz | EN 302 537 |
| b | 9 -315 KHz | EN 302 195 |
| c | 316 – 600KHz | EN 302 536 |
| d | 30 – 37.5 MHz | EN 302 510 |
| e | 12.5 – 20 MHz | EN 300 330 |

#### General Non- Specific SRD’s

The designations used in Table 5.5 below (a to m) are cross reference with those used in table 5.3 above.

Table 5.5 Harmonized standards for non specific SRD’s

|  | Frequency | ETSI Standard |
| --- | --- | --- |
| a | 6765-6795 kHz | EN 300 330 |
| b | 13.553-13.567 MHz | EN 300 330 |
| c | 26.957-27.283 MHz | EN 300 220 |
| d | 40.660-40.700 MHz | EN 300 220 |
| e | 138.20-138.45 MHz | EN 300 220 |
| f | 433.050-434.790 MHz ( see note 4) | EN 300 220 |
| f1 | 433.050-434.790 MHz (note 4bis) | EN 300 220 |
| f2 | 434.040-434.790 MHz (note 4bis) | EN 300 220 |
| g | 863-870 MHz (note 3, 4 and 6) | EN 300 220 |
|  |  | EN 300 220 |
|  |  | EN 300 220 |
| g1 | 868.000-868.600 MHz (note 4) | EN 300 220 |
| g2 | 868.700-869.200 MHz (note 4) | EN 300 220 |
| g3 | 869.400-869.650 MHz (note 4) | EN 300 220 |
| g4 | 869.700-870.000 MHz (note 4bis) | EN 300 220 |
| h | 2400.0-2483.5 MHz | EN 300 440 |
| i | 5725-5875 MHz | EN 300 440 |
| j | 24.00-24.25 GHz | EN 300 440 |
| k | 61.0-61.5 GHz | N/S |
| l | 122-123 GHz | N/S |

## UWB Frequency Band

### Inside CEPT

In the CEPT (45 countries involved), two decisions were taken.

The Decision ECC/DEC/(06)12 amended 31 October 2008 supplements ECC/DEC/(06)04. The latter was taken the 24 of March 2006 allows UWB operation in the upper band (6 to 8.5 GHz ) with the following emission limits (Table below), while the former was taken into effect from 1 May 2009. This decision as now to be implemented by national regulatory agencies of the 45 CEPT member countries.

The main points of this decision are :

- that this ECC Decision defines general harmonised conditions for the use in Europe of devices using UWB technology in bands below 10.6 GHz;

- that the devices permitted under this ECC Decision are exempt from individual licensing and operate on a non-interference, non-protected basis;

- that this ECC Decision is not applicable to:

  - flying models such as R/C models

  - outdoor installations and infrastructure, including those with externally mounted antennas,

  - devices installed in road and rail vehicles, aircraft and other aviation;

- that devices covered by the scope of this ECC Decision are not allowed to be used at a fixed outdoor location or connected to a fixed outdoor antenna;

- that the technical requirements detailed in Annex 1 apply to devices permitted under this ECC Decision;

- that this Decision enters into force on 24 March 2006;

- that the preferred date for implementation of this Decision shall be 1 October 2006;

- that CEPT administrations shall communicate the national measures implementing this Decision to the ECC Chairman and the Office when the Decision is nationally implemented.”

ECC Decision 06/04 on UWB emission limits, EIRP in dBm/MHz

| Frequency range | Maximum mean e.i.r.p. density (dBm/MHz) | Maximum peak e.i.r.p. density (dBm/50MHz) (Note 2) |
| --- | --- | --- |
| Below 1.6 GHz | –90 dBm/MHz | -50 dBm/50MHz |
| 1.6 to 3.8 GHz (Note 1) | -85 dBm/MHz | –45 dBm/50MHz |
| 3.8 to 4.8 GHz (Note 1) | –70 dBm/MHz | –30 dBm/50MHz |
| 4.8 to 6 GHz | –70 dBm/MHz | –30 dBm/50MHz |
| 6 to 8.5 GHz | –41.3 dBm/MHz | 0 dBm/50MHz |
| 8.5 to 10.6 GHz | –65 dBm/MHz | -25 dBm/50MHz |
| Above 10.6 GHz | –85 dBm/MHz | -45 dBm/50MHz |

Note 1: 	ECC is still considering whether or not to adopt a separate Decision covering the frequency band 3.1 – 4.8 GHz.

Note 2: 	The peak e.i.r.p. can be alternatively measured in a 3 MHz bandwidth. In this case, the maximum peak e.i.r.p. limits to be applied is scaled down by a factor of 20log(50/3) = 24.4 dB.

A separate decision (ECC/DEC(06)12) has been taken concerning "low duty cycle" limitation to enable access to the use of the lower band with the guarantee of an efficient protection of licensed services : To permit uses of the low band (3.4 GHz - 4.8 GHz) to low activity applications for which this band is essential, the European regulator as also defined a mitigation technique called "Low Duty Cycle" (LDC). A device implementing LDC is an UWB device as stated under the generic rules who also meets the following requirements:

Ton max = 5 ms

Toff mean ≥ 38 ms (averaged over 1 sec)

Σ Toff > 950 ms per second

Σ Ton < 5% per second and 0.5% per hour

- Ton

Ton is defined as the duration of a burst irrespective of the number of pulses contained.

- Toff

Toff is defined as the time interval between two consecutive bursts when the UWB emission is kept idle.

Supplementary regulatory provisions from ECC/DEC/(06)04 are as follows.

- that this ECC Decision defines supplementary regulatory provisions to Decision ECC/DEC/(06)04 for UWB devices using mitigation techniques;

- that, for the purpose of the Decision, the following definitions apply:

  - LDC UWB devices are defined as devices using UWB technology that meet the technical requirements for Low Duty Cycle (LDC) mitigation technique given in Annex 1 of this Decision;

  - DAA UWB devices are defined as devices using UWB technology that meet the technical requirements for Detect And Avoid (DAA) mitigation technique given in Annex 2 of this Decision;

- that the devices permitted under this ECC Decision are exempt from individual licensing and operate on a non-interference, non-protected basis;

- that this ECC Decision is not applicable to:

  - devices and infrastructure installed at a fixed outdoor location or devices connected to a fixed outdoor antenna;

  - devices installed in flying models, aircraft and other aviation;

- that for devices permitted under this ECC Decision, the technical requirements detailed in Annex 1 of Decision ECC/DEC/(06)04 apply, and, alternatively:

  - within the band 3.1 – 4.8 GHz, LDC UWB devices are permitted to operate with a maximum mean e.i.r.p. spectral density of -41.3 dBm/MHz and a maximum peak e.i.r.p. of 0 dBm measured in 50MHz; notwithstanding that

    - in case of LDC UWB devices installed in road and rail vehicles, no additional mitigation is required as the LDC mitigation technique is recognised to offer a protection level that is at least equivalent to Transmit Power Control (TPC);

  - within the bands 3.1 – 4.8 GHz and 8.5 – 9 GHz, DAA UWB devices are permitted to operate with a maximum mean e.i.r.p. spectral density of -41.3 dBm/MHz and a maximum peak e.i.r.p. of 0 dBm measured in 50MHz; notwithstanding that,

    - in case of DAA UWB devices installed in road and rail vehicles, operation is subject to the implementation of Transmit Power Control (TPC) with a range of 12 dB with respect to the maximum permitted radiated power. If no TPC is implemented, the maximum mean e.i.r.p. spectral density is -53.3 dBm/MHz;

- that this Decision enters into force on 31 October 2008;

- that the preferred date for implementation of this Decision shall be 1 May 2009;

that CEPT administrations shall communicate the national measures implementing this Decision to the ECC Chairman and the Office when the Decision is nationally implemented.”

### Inside the European Union (27 countries who are also CEPT Member)

Within the European Union, the Radio Spectrum Comittee (RSC) has the authority to harmonize spectrum use. RSC has decided to enforce ECC Decision and also to make it less restrictive. It should be noted that every EU-member is also a CEPT-Member. In those countries, RSC rules are applied. Therefore, the RSC decision will be applicable throughout the whole of the European Union.

For the purposes of this Decision, definition are :

1. “equipment using ultra-wideband technology” means equipment incorporating, as an integral part or as an accessory, technology for short-range radiocommunication, involving the intentional generation and transmission of radio-frequency energy that spreads over a frequency range wider than 50 MHz, which may overlap several frequency bands allocated to radio communication services;

2. “non-interference and non-protected basis” means that no harmful interference may be caused to any radiocommunication service and that no claim may be made for protection of these devices against harmful interference originating from radiocommunication services;

3. “indoors” means inside buildings or places in which the shielding will typically provide the necessary attenuation to protect radiocommunication services against harmful interference;

4. "automotive vehicle" means any vehicle as defined by Article 1 of Directive 70/156/EEC;

5. "railway vehicle" means any vehicle as defined by Article 1 of (EC) Regulation 1192/2003;

6. “e.i.r.p.” means Equivalent or Effective Isotropic Radiated Power;

7. “mean e.i.r.p. density” means the mean power measured with a 1 MHz resolution bandwidth, a root-mean-square (RMS) detector and an averaging time of 1 ms or less;

8. “peak e.i.r.p. density” means the peak level of transmission contained within a 50 MHz bandwidth centred on the frequency at which the highest mean radiated power occurs. If measured in a bandwidth of x MHz, this level is to be scaled down by a factor of 20log(50/x)dB;

“maximum e.i.r.p. density” means the highest signal strength measured in any direction at any frequency within the defined range; and limitations are given in the following tables (2009/343/EC)

- “building material analysis” (BMA) means a field disturbance sensor that is designed to detect the location of objects within a building structure or to determine the physical properties of a building material;

- “radiated into the air” means those parts of the signal emitted by specific applications of ultra-wideband tech-nology which are not absorbed by their shielding or by the material under investigation.’.

Equipment using ultra-wideband technology may also be allowed to use the radio spectrum with other e.i.r.p. limits together with other appropriate mitigation techniques on condition that it achieves at least an equivalent level of protection as provided by the limits in this table.

A. Appropriate mitigation techniques

Equipment using ultra-wideband technology shall also be allowed to use the radio spectrum with higher e.i.r.p. limits than mentioned in the table in section 1.1 when applying additional mitigation techniques as described in the relevant harmonised standards adopted under Directive 1999/5/EC or other mitigation techniques on condition that it achieves at least an equivalent level of protection as provided by the limits in the table in section 1.1. The following mitigation techniques are presumed to provide such protection:

A.1. “Low duty cycle” (LDC) mitigation

A maximum mean e.i.r.p. density of – 41,3 dBm/MHz and a maximum peak e.i.r.p. of 0 dBm measured in 50 MHz shall be allowed in the 3,1-4,8 GHz bands provided that a low duty cycle restriction is applied in which the sum of all transmitted signals is less than 5 % of the time each second and less than 0,5 % of the time each hour, and provided that each transmitted signal does not exceed 5 ms.

A.2. “Detect and avoid” (DAA) mitigation

A maximum mean e.i.r.p. density of – 41,3 dBm/MHz and a maximum peak e.i.r.p. of 0 dBm measured in 50 MHz shall be allowed in the 3,1-4,8 GHz and 8,5-9,0 GHz bands provided that a detect and avoid (DAA) mitigation technique as described in the relevant harmonised standards adopted under Directive 1999/5/EC is used.

Main changes compared to ECC decisions are :

- Simplification of the Low Duty Cycle criteria

Allow mean EIRP up to -41.3 dBm/MHz between 4.2 - 4.8 GHz until 12/31/2010

### Relevant ETSI Harmonised Standards

For general UWB communication devices EN 302 065 is correct Standard. This standard is currently being reviewed in order to insert test methods and limits for Detect And Avoid (DAA) requirements. This work is being carried out in ETSI ERM TG31A.

There are other UWB application specific Standards as follows and activities related to these standards take place in TG31B and TG31C:

EN 302 288 - 24 GHz Short Range Radar Harmonised Standard

EN 302 264 - Ultra Wide Band Radar Equipment Operating above 60 GHz

EN 302 498 – Object Discrimination and Characterization

EN 302 499 – Object Identification for Surveillance

EN 302 501 – Ultra Wide Band (UWB) Sensors for Location Tracking type 2

EN 302 619 – Object Detection Mobile in the 6-7.25 GHz range

# Japan

This section describes the frequency regulation for medical application in Japan. Frequency regulations are usually regulated by Ministry of Inner Affairs and Communications (MIC) in Japan.

## Medical Frequency Bands

### Wireless Medical Telemetry System (WMTS)

Band plan for WMTS in Japan is as follows.

### Medical Implant Communication Systems (MICS)

This section is an English translation of the summary of a report of Telecommunications Concil issued by MIC on February 28, 2005.

Definition:

MICS is a system that conducts data transmission between a medical implant device (MID), which can be an implanted device or a device temporarily placed in a body, and a medical radio device outside body (programme device), which is used to control the MID, for the purpose of data transmission to start, change, or terminate a function of vital signal or a device.

Main Technical Conditions for Radio Device:

|  | Main Technical Conditions for Radio Device |
| --- | --- |
| Frequency bands | From 402 MHz to 405 MHz (secondary use*1) |
| Communication method | Simplex, one-way, or duplex. |
| Radio Types | Two-sided Amplitude Modulation, Frequency Modulation, or Phase Modulation. Use only single carrier modulation. |
| Bandwidth occupancy | Smaller than or equal to 300 kHz |
| Frequency tolerance | Within +/- 100*10-6 |
| Transmission power | Smaller than or equal to 25uW in EIRP*2 |
| Transmission power tolerance | Upper limit is 20% and lower limit is -50% |
| Out-band and spurious emission level | - Permitted out-band or spurious emission level It shall be smaller or equal to 25nW in EIRP*2 at frequencies 150 kHz lower or beyond the centre frequency. - Reference bandwidth Unwanted emission frequencies; Reference bandwidth. 9 kHz < f < 150 kHz; 1 kHz. 150 kHz < = f < 30 MHz; 10 kHz. 30 MHz < = f < 1 GHz; 100 kHz. f > 1GHz; 1 MHz. |
| Function to avoid cross intercession | The programme device shall have the function of channel sensing before it starts to send signal. The MID is not required to have this function. |

*1 MICS shall not cause interference to the first order service of meteorology supporting radio and allow the interference from the latter.

*2 For MID, this is defined as the emission from the body surface.

## UWB Frequency Band

### Spectrum Mask

The spectrum mask regulated by the Telecommunication Council of Ministry of Internal Affairs and Communications (MIC) is shown in Figure 6.1.

Figure 6.1 Japanese spectrum mask (only indoor use).

The frequency bands of 3400 MHz through 4800 MHz and 7250 MHz through 10250 MHz are assigned for UWB operation. For 3400 MHz through 4800 MHz, interference mitigation techniques are required. However, for 4200 MHz through 4800 MHz, interference mitigation techniques are not required until the end of December, 2010. UWB systems shall not interrupt other radio systems operated in the same band. UWB systems can not defer the operation of other radio systems.

### General Technical Requirements

This part is according to Article 4.4 of Regulations for Enforcement of Radio Law issued by MIC Japan. It should be noted the Japanese regulation ask that at least one of the UWB radio equipments in a group such as a piconet shall be connected to the AC main power supply. In more detail, if a UWB radio equipment is connected to the AC mains power supply, it can start transmission at any time. A UWB radio equipment not connected to the AC mains power supply shall be permitted to emit radio waves only after it receives a signal from another radio equipment connected to the AC mains power supply.

#### General Conditions

###### Contents of communications

Content of communication shall primarily be for data transmissions.

###### Communication methods

Communication methods shall be either simplex operation, full duplex operation or half duplex operation.

###### Operating Frequency band

Operating frequency band to be used shall be from 3.4 GHz to 4.8 GHz or from 7.25 GHz to 10.25 GHz.

###### Usage environment condition

The radio equipment shall be used indoors.

#### Transmitter

- Modulation method

There is no specific restriction on the modulation method

- Antenna power

Antenna power for the operating frequency band shall satisfy respective values as follows:

- Average power per 1MHz bandwidth shall be -41.3 dBm or less.

As an interim measure until December 31, 2010, the following condition shall apply: Of radio equipment with use of frequencies from 3.4 GHz to 4.8 GHz, those without an interference mitigation function stipulated in 3.4 (2) shall have the average power of -70 dBm or less per 1 MHz bandwidth within the frequency band from 3.4 GHz to 4.2 GHz, and the average power of -41.3 dBm or less per 1 MHz bandwidth within the frequency band from 4.2 GHz to 4.8 GHz.

- Peak power per 50MHz bandwidth shall be 0 dBm or less.

As an interim measure until December 31, 2010, the following condition shall apply: Of radio equipment with use of frequencies from 3.4 GHz to 4.8 GHz, those without an interference mitigation function stipulated in 3.4 (2) shall have a peak power of -30 dBm or less per 50 MHz bandwidth within the frequency band from 3.4 GHz to 4.2 GHz, and the peak power of 0 dBm or less per 50 MHz bandwidth within the frequency band from 4.2 GHz to 4.8 GHz.

- Tolerances for antenna power

The tolerance of the antenna power variation shall be less than the allowed maximum value + 20%.

- Transmission data rate

Transmission data rate shall be higher than 50 Mbps excepting for such cases as noise or interference from other radio stations need be avoided.

- Frequency bandwidth

Frequency bandwidth between the upper and lower frequencies for which the radiation power drops 10 dB below the maximum radiation power shall be 450 MHz or more.

- Permissible values for occupied bandwidth

Permissible values for the occupied bandwidth shall be as follows:

For those with use of the frequency from 3.4 GHz to 4.8 GHz		1.4 GHz

For those with use of the frequency from 7.25 GHz to 10.25 GHz	3 GHz

- Permissible values for unwanted emission intensity

Permissible values for the unwanted emission intensity from the transmitter are as follows:

Note: In the case that the radio spectrum is located within a frequency range from 3.4 GHz to 4.8 GHz, the antenna power for the frequency band from 3.4 GHz to 4.8 GHz shall comply with Article 3.2 (2) instead of 3.2 (7). Similarly, in the case that the radio spectrum is located within a frequency range from 7.25 GHz to 10.25 GHz, the antenna power for the frequency band from 7.25 GHz to 10.25 GHz shall comply with the Article 3.2 (2) instead of 3.2 (7).

Provision by peak power

Frequency [GHz]

Power density [dBm / MHz]

1)  With an interference mitigation function

Antenna power density

Permissible values for unwanted emission intensity

Provision by average power

Power density [dBm]

MHz]

Antenna power density

Permissible values for unwanted emission intensity

Provision by average power

Provision by peak power

Frequency [GHz]

2)  Without an interference mitigation function (until December 31, 2010)

Figure 6.2 Permissible values for unwanted emission intensity 
(when using frequency from 3.4 GHz to 4.8 GHz)

Power density [dBm]

MHz]

Antenna power density

Frequency [GHz]

Permissible values for

unwanted emission intensity

Provision by average power

Provision by peak power

Figure 6.3　 Permissible values for unwanted emission intensity 
(when using frequency from 7.25 GHz to 10.25 GHz)

#### Receiver

Limit on Secondary Radiated Emissions, etc.

| Frequency band | Average power at any 1 MHz bandwidth |  |
| --- | --- | --- |
|  | When using frequency from 3.4 GHz to 4.8 GHz | When using frequency from 7.25 GHz to 10.25 GHz |
| Less than 1,600 MHz | -90 dBm or less |  |
| From 1,600 MHz to 2,700 MHz | -85 dBm or less |  |
| From 2,700 MHz to 3.4 GHz | -70 dBm or less |  |
| From 3.4 GHz to 4.8 GHz | -54 dBm or less | -70 dBm or less |
| From 4.8 GHz to 7.25 GHz | -70 dBm or less |  |
| From 7.25 GHz to 10.25 GHz | -70 dBm or less | -54 dBm or less |
| From 10.25 GHz to 10.6 GHz | -70 dBm or less |  |
| From 10.6 GHz to 10.7 GHz | -85 dBm or less |  |
| From 10.7 GHz to 11.7 GHz | -70 dBm or less |  |
| From 11.7 GHz to 12.75 GHz | -85 dBm or less |  |
| More than 12.75 GHz | -70 dBm or less |  |

Power density [dBm]

MHz]

Antenna power density

Frequency [GHz]

1)  With an interference mitigation function

Power density [dBm] MHz]

Antenna power density

Frequency [GHz]

2)　 Without an interference mitigation function (until December 31, 2010)

Figure 6.4　Limit on Secondary Radiated Emissions 
(when using frequency from 3.4 GHz to 4.8 GHz)

Power density [dBm]

MHz]

Antenna power density

Frequency [GHz]

Figure 4.5　Limit on Secondary Radiated Emissions 
(when using frequency from 7.25 GHz to 10.25 GHz)

#### Controller

A controller shall have the following devices and functions, and meet each of the following conditions.

###### Interference prevention function

The radio equipment of the radio station shall mainly be used in the same premises. It shall transmit automatically identification signs or receive them.

###### Interference mitigation function

The radio equipment using radio with frequency from 3.4 GHz to 4.8 GHz shall have the interference mitigation function that meets the technical conditions notified separately by Minister of Internal Affairs and Communications.

As an interim measure until December 31, 2010, it is not mandatory to install an interference mitigation function. When the interim measure is applied, the description of 3.2 (2) shall be noted for the antenna power, the description of 3.2 (7) for permissible values for unwanted emission intensity and the description of 3.3 (1) for the limit on secondary radiated emissions, etc.

###### Controls of radio emission

A radio equipment not connected to the AC mains power supply shall be permitted to emit radio waves only after it receives a signal from another radio equipment connected to the AC mains power supply.

#### Connection with Telecommunication Circuit Equipment

A radio equipment that is connected to the telecommunication circuit equipment shall meet the following conditions:

###### Identification devices

An identification device shall have a unique identification sign (sign designed to identify the radio equipment used for the terminal equipment and to be collated when configuring a communication channel).

###### Identification signs

The coding length of the identification sign shall be 48 bits or more.

###### Method to judge the availability of the operating frequency band

Judgment of the availability of the operating frequency band shall be made by detecting a radio wave emitted from another radio station, or by calculating the received signal and detecting the signal level. However, for the equipment that has a function to disconnect the channel when the communication quality is degraded, the judgment may be replaced by checking the operating condition of the channel.

#### Antenna

###### Antenna structures

There is no specific provision for the antenna structure.

###### Gain of the transmitting antenna

Absolute gain of the transmitting antenna shall be 0 dBi or less. However, in the case that the e.i.r.p. (Equivalent Isotropic Radiated Power) falls short of the antenna power provided in 3.2 (2) added with the transmitting antenna gain of 0 dBi absolute gain, the difference may be complemented by adjusting the gain of the transmitting antenna.

###### Gain of the receiving antenna

There is no specific provision for the gain of the receiving antenna.

#### Others

###### Cabinet

The cabinet shall be constructed so as to be not be easily tampered with.

###### Mark of usage restrictions

A mark of the usage restriction that radio waves can be emitted only indoor shall be visibly shown on the cabinet.

###### Mark in relation to technical regulations conformity certification

A mark in relation to technical regulations conformity certification in the specified format shall be visibly shown on the radio equipment.

###### Mark in relation to technical conditions compliance approval for terminal equipment

###### In case a radio equipment connects to the telecommunication circuit equipment, a mark in relation to technical conditions compliance approval for terminal equipment in the specified format shall be visibly shown on it.

# Korea

This section describes the frequency regulation for medical applications in Korea.

## Medical Frequency Bands

(The sources)

- The notice 2007-35 of MIC (Revision 2007.10.19.)

Wireless system for radio stations where can be opened without report

* MIC : Ministry of Information and Communication

- The notice 2007-80 of RRL (Revision 2007.10.17.)

Technical standards for wireless equipment for businesses except for broadcasting, marine, aviation, and telecommunications.

* RRL : Radio Research Laboratory(A government agency directly controlled by MIC)

1. A definition of MICS(Medical Implant Communication System) in the notice 2007-35 is as follows ;

Wireless systems for MICS are composed of wireless equipment implanted into bodies to diagnosis and to treat a patient and wireless equipment placed out-body to control.

Wireless system for MICS is as follows ;

| Frequency Band(㎒) | Radiation Power | Remarks |
| --- | --- | --- |
| 402～405 | Not exceeding 25㎼ (Including antenna gain) | Not exceeding 9.1㎶/m @ 3m |

2. Technical standards using 402 ~ 405 MHz frequency band conform with the following provisions ;

(1) Besides emergency, an implantable wireless system shall communicate with wireless control system only which is placed in out-body

(2) A radiation power including antenna gain shall not exceed 25 ㎼.

(3) A frequency bandwidth shall not exceed 300kHz(BW at point of -20dB to maximum power).

(4) A number of frequency channel shall be more than 9 and each channel shall not be overlapped respectively.

(5) Allowable frequency variation falls below ±100×10-6 at -20○C  ~ 50○C for wireless system and at -20○C  ~ 45○C for implantable wireless system.

(6) An unwanted radiation in spurious area shall not exceed the following criteria

| Frequency | Under 1 ㎓ | Above 1 ㎓ |
| --- | --- | --- |
| Criteria | - 36 ㏈m | - 30 ㏈m |

(7) A wireless control system shall have channel-selectable function to establish communication channels before starting communication with wireless implantable system.

| Items | Reference |
| --- | --- |
| Interference sensing reference | - 10logB(Hz)-150(dBm/Hz)+G(dBi) Here, B : maximum bandwidth on communication (max BW among BWs at point of -20dB to Max power) G : absolute gain of wireless control system |
| Time for checking a receiving power per channel | Above 10 msec |
| Time for checking the available channel and initiating communication | Not exceeding 5 sec |

a) Wireless control systems shall start to communicate within 5 seconds after assuring a communication channel those receiving power is lower than interference sensing reference and an extra channel ready for interference. Where, a channel those receiving power is the lowest one can be selected to communicate in case that all channels are exceed the receiving power reference.

b) The following functions are possessed to switch into a spare channel in a   circumstance that a normal communication is unavailable because of interference during communication.

o In case that a receiving power is re-measured before switching to a spare channel and its value is not higher than the initial value by 6dB : To communicate by switching to the pre-set spare channel

o In case that a receiving power is re-measured before switching to a spare channel and its value is higher than the initial value by 6dB : To communicate by repeating the initial process of channel setting.

(8) Establish a function to automatically stop communication if there is no transmission and receipt of a data in 5 seconds after starting communication.

(9) The above (1), (2), (4), and (7) are not applied to a wireless implantable system which uses 1 channel in 403.5 ~ 403.8MHz band and its output power is lower than 100nW(e.i.r.p.).

## UWB Frequency Band

1. Wireless system using UWB technology is as follows ;

| Frequency Band | Radiation Power | Remarks |
| --- | --- | --- |
| 3.1～4.8 7.2～10.2 | Not exceeding 41.3dBm/㎒ (Including antenna gain) | Confine system using UWB technology |
| 57～64 | Not exceeding 10 mW |  |

2. Technical standards using UWB technology shall conform with the following provisions ;

(1) Frequency Band and Power density

| Frequency Band(㎓) | Power density including antenna gain |  | Remarks |
| --- | --- | --- | --- |
|  | Mean Power | Maximum Power |  |
| 3.1～4.8 7.2～10.2 | -41.3 ㏈m/㎒ | 0 ㏈m/50 ㎒ | Power density shall conform with both mean power and maximum power |

(2) General Conditions : Do not apply to aeroplanes, ships, satellites and R/C models.

(3) Antenna and wireless system shall be one body :

(4) Frequency Bandwidth(BW which is -10dB from Maximum power) shall be more than 450MHz

(5) Undesirable radiation power conform with the following conditions;

| Frequency Band | Mean power density including Antenna gain |
| --- | --- |
| Lower than 1.6 ㎓ | Not exceeding -90 ㏈m/㎒ |
| Above 1.6 ㎓to lower than 2.7 ㎓ | Not exceeding -85 ㏈m/㎒ |
| Above 2.7 ㎓to lower than 3.1 ㎓ | Not exceeding -70 ㏈m/㎒ |
| Above 3.1 ㎓to lower than 4.8 ㎓ | Not exceeding -51.3 ㏈m/㎒ |
| Above 4.8 ㎓to lower than 7.2 ㎓ | Not exceeding -70 ㏈m/㎒ |
| Above 7.2 ㎓to lower than 10.2 ㎓ | Not exceeding -51.3 ㏈m/㎒ |
| Above 10.2 ㎓ | Not exceeding -70 ㏈m/㎒ |

(6) Wireless systems using 3.1～4.8 ㎓ frequency band conform with one of the following DAA or LDC conditions;

a) Mean power density including Antenna gain shall not exceed  -70 ㏈m/㎒

b) Continuous transmission time shall not exceed 5 msec and recess time shall be more 1 sec

c) If it receives a signal from other stations more -80 ㏈m/㎒ level during operation, the wireless system shall decrease level of its signal not exceeding -70 ㏈m/㎒ within 2 sec

d) If it receives a signal from other stations more -80 ㏈m/㎒ level during operation, the wireless system shall avoid a collision within 2 sec

(7) Secondary Emission shall not exceed -54 ㏈m/㎒ in using frequency band.

# UK

This section describes the relevant information on the regulations of UK pertaining to BAN usage.

## General Non-Specific Short Range Devices

Minimum requirements for the use of general non-specific short range devices (SRD) are described in this clause.

Wireless telegraphy apparatus designed or adapted for Non-Specific SRD, so as to be capable of use on one or more of the frequencies or within one of the frequency bands, and at a radiated level not exceeding the maximum for such Frequency / Bands, for each category of apparatus, specified in the table below and subject to the following sub-paragraphs:

(a) For category i and ii, voice is permitted using frequency modulation only.

(b) In category vi, channel numbers 1, 3 - 11; are available with a channel centre frequency of 173.2 MHz + (Channel B/width x channel number).

(c) In category vii, channel numbers 2 - 5 are available with a channel centre frequency of 173.2 MHz + (Channel B/width x channel number).

(d) In category viii, telemetry and telecommand may only be used in conjunction with telephony with a non-locking push to talk key or voice operated carrier.

(e) In category ix, the band 417.9 - 418.1 MHz will be withdrawn for this application after the 31 December 2007. Equipment placed on the market under the R&TTE Directive on or before 31 December 2002 may be taken into service until 31 December 2007. After 31 December 2007, equipment can no longer be taken into service. However existing equipment may continue to operate within the band provided that interference is not caused to the licensed service.

(f) In category xii, the band may also be used for airborne telemetry applications with channel spacing of 25 kHz.

(g) In categories xiii, xiv, xvi and xvii, consecutive channels may be combined where a larger bandwidth is required, due to the modulation of the signal, up to the maximum sub-band frequency allocation. The total signal bandwidth, including frequency error and drift, must be contained within the allocated sub-band under all conditions. Where a duty cycle limit is applied, that duty cycle limit remains in force where channels are combined.

(h) In category xv, a channel access protocol shall be used.

(i) For categories v, vi, vii, ix, xi, xii & xvii, Music and Speech is permitted when using a digitised signal.

| category | Frequency / Band | Radiated Level | Channel b/width | Music / Speech | Duty cycle | Reference Standard * |
| --- | --- | --- | --- | --- | --- | --- |
| i | 26.995, 27.045, 27.095, 27.145, 27.195 MHz | 10 mW e.r.p. | 10 kHz | Yes | - | EN 300 220 |
| ii | 40.66 - 40.70 MHz | 10 mW e.r.p. | 10 kHz | Yes | - |  |
| iii | 49.82 - 49.98 MHz | 10 mW e.r.p. | 10 kHz | Yes | - |  |
| iv | 49.82 - 49.98 MHz | 10 mW e.r.p. | - | Yes | - |  |
| v | 40.66 - 40.7 MHz | 10 mW e.r.p. | - | Yes | - |  |
| vi | 173.20 - 173.35 MHz | 1 mW e.r.p. | 12.5 kHz | Yes | - |  |
| vii | 173.20- 173.35 MHz | 1 mW e.r.p. | 25 kHz | Yes | - |  |
| viii | 173.5875, 173.6 MHz | 10 mW e.r.p. | 12.5 kHz | Yes | - |  |
| ix | 417.9 - 418.1 MHz | 250 μW e.r.p. | - | Yes | - |  |
| x | 433.05- 434.79 MHz | 10 mW e.r.p. | - | Yes | ≤10% |  |
| xi | 433.05- 434.79 MHz | 1 mW e.r.p. | - | Yes | - |  |
| xii | 434.04- 434.79 MHz | 10 mW e.r.p. | 25 kHz | Yes | - |  |
| xiii | 868.0- 868.6 MHz | 25 mW e.r.p. |  | Yes | ≤ 1 % |  |
| xiv | 868.7- 869.2 MHz | 25 mW e.r.p. |  | Yes | ≤ 0.1 % |  |
| xv | 869.30- 869.40 MHz | 10 mW e.r.p. | ≤ 25 kHz | Yes | ≤ 10 % |  |
| xvi | 869.40- 869.65 MHz | 500 mW e.r.p. | ≤ 25 kHz | Yes | ≤10 % |  |
| xvii | 869.70- 870.00 MHz | 5 mW e.r.p. |  | Yes | - |  |
| xviii | 2400- 2483.5 MHz | 10 mW e.i.r.p. | - | Yes | - | EN 300 440 |
| xix | 5725- 5875 MHz | 25 mW e.i.r.p. | - | Yes |  |  |

## Industrial / Commercial Telemetry and Telecommand

Minimum requirements for the use of industrial / commercial telemetry and telecommand are described in this clause.

Wireless telegraphy apparatus designed or adapted for Telemetry and Telecommand, so as to be capable of use on one or more of the frequencies or within one of the frequency bands, and at a radiated level not exceeding the maximum for such Frequency / Bands, for each category of apparatus, specified in the table below and subject to the following sub-paragraphs:

(a) In category i, channel numbers 1, 3 to 11; are available with a channel centre frequency of 173.2 MHz + (Channel B/width x channel number).

(b) In category ii, channel numbers 1 to 5 are available with a channel centre frequency of 173.2 MHz + (Channel B/width x channel number).

(c) In category iv, channel numbers 1 to 25, 28 to 31 and 33 to 35 are available with a channel centre frequency of 458.5 MHz + (Channel B/width x channel number).

(d) In category v, channel numbers 1 to 12, 14 to 15 and 17 are available with a channel centre frequency of 458.5 MHz + (Channel B/width x channel number).

(e) In category vi, consecutive channels may be combined where a larger bandwidth is required, due to the modulation of the signal, up to the maximum sub-band frequency allocation. The total signal bandwidth, including frequency error and drift, must be contained within the allocated sub-band under all conditions. Where a duty cycle limit is applied, that duty cycle limit remains in force where channels are combined.

(f) In category viii the use is limited to remote meter reading and in category ix the use is limited to Asset Tracking and Tracing

(g) For categories i, ii, iii, iv, v & vii, Music and Speech is permitted when using a digitised signal.

| category | Frequency / Band | Radiated Level | Channel B/width | Music / Speech | Duty cycle | Reference Standard * |
| --- | --- | --- | --- | --- | --- | --- |
| i | 173.2 - 173.35 MHz | 10 mW e.r.p. | 12.5 kHz | Yes | - | EN 300 220 |
| ii | 173.2 - 173.35 MHz | 10 mW e.r.p. | 25 kHz | Yes | - |  |
| iii | 173.2 - 173.35 MHz | 10 mW e.r.p. | - | Yes | - |  |
| iv | 458.5 - 458.95 MHz | 500 mW e.r.p. | 12.5 kHz | Yes | - |  |
| v | 458.5 - 458.95 MHz | 500 mW e.r.p. | 25 kHz | Yes | - |  |
| vi | 869.40 to 869.65 MHz | 500 mW e.r.p. | ≤25 kHz | Yes | ≤10% |  |
| vii | 2445 - 2455 MHz | 100 mW e.i.r.p. | - | Yes | - | EN 300 440 |
| viii | 169.4 – 169.475 MHz | 10mW e.r.p. | 50kHz | Yes | <10% | EN 300 220 |
| ix | 169.4 – 169.475 MHz | 10mW e.r.p. | 50kHz | Yes | <1% | EN 300 220 |

## Medical and Biological Applications

Minimum requirements for the use of medical and biological applications are described in this clause.

Wireless telegraphy apparatus designed or adapted for Telemetry and Telecommand, so as to be capable of use on one or more of the frequencies or within one of the frequency bands, and at a radiated level not exceeding the maximum for such Frequency / Bands, for each category of apparatus, specified in the table below and subject to the following sub-paragraphs:

(a) In category iii, channel numbers 1 to 24 are available with channel centre frequency of 173.7 MHz + (Channel B/width x channel number).

(b) In category iv, channel numbers 1 to 11 are available with channel centre frequency of 173.7 MHz + (Channel B/width x channel number).

(c) In category vi, for use with ultra low power active medical implants only. Individual transmitters may combine adjacent channels for increased bandwidth with advanced mitigation techniques.

(d) In category vii and viii, channel numbers 37 to 47 are available with channel centre frequency of 458.5 MHz + (Channel B/width x channel number).

(e) In category ix and x, channel numbers 19 to 23 are available with channel centre frequency of 458.5 MHz + (Channel B/width x channel number).

(f) In categories iii, iv, vii and ix, these bands may also be used for the tracking of birds.

| category | Frequency / Band | Radiated Level | Channel B/width | Music / Speech | Duty cycle | Reference Standard * |
| --- | --- | --- | --- | --- | --- | --- |
| i | 300 kHz–- 30 MHz | 9 dBμA/m @10 m | - | - | - | EN 300 330 |
| ii | 30 to 37.5 MHz | 1 mW e.r.p. | - | - | 10% | EN 300 220 |
| iii | 173.7 – 174 MHz | 10 mW e.r.p. | 12.5 kHz | - | - |  |
| iv | 173.7 – 174 MHz | 10 mW e.r.p. | 25 kHz | - | - |  |
| v | 173.7 – 174 MHz | 10 mW e.r.p. | - | - | - |  |
| vi | 402 – 405 MHz | 25 μW e.r.p. | 25 kHz | - | - | EN 301 839 |
| vii | 458.9625 – 459.1000 MHz | 10 mW e.r.p. | 12.5 kHz | - | - | EN 300 220 |
| viii | 458.9625 – 459.1000 MHz | 500 mW e.r.p. | 12.5 kHz | - | - |  |
| ix | 458.9625 – 459.1000 MHz | 10 mW e.r.p. | 25 kHz | - | - |  |
| x | 458.9625 - 459.1000 MHz | 500 mW e.r.p. | 25 kHz | - | - |  |
| xi | 9 – 135 kHz | 30 dBμA/m @ 10m | - | - | 10 % | EN 302 195 |

# USA

This section describes the frequency regulation for medical applications in USA. Frequency regulations are regulated by the Federal Communications Commission (FCC).

## Medical Frequency Bands

### Wireless Medical Telemetry System (WMTS)

Wireless medical telemetry generally is the remote monitoring of a patient's health through radio technology. The use of wireless medical telemetry gives patients greater mobility and increased comfort by freeing them from the need to be connected to hospital equipment that would otherwise be required to monitor their condition. Wireless medical telemetry also serves the goal of reducing health care costs because it permits the remote monitoring of several patients simultaneously.

All types of communications except voice and video are permitted on both a bi-directional and unidirectional basis, provided that all communications are related to the provision of medical care. The FCC rules and regulations are codified in Subpart H of Title 47 of the Code of Federal Regulations (CFR) 95.1101-95.1129 and are summarized in the following sections.

#### WMTS Transmitter frequencies

WMTS transmitters may operate in the frequency bands specified as follows:

608–614 MHz

1395–1400 MHz

1427–1429.5 MHz except at the locations listed in CFR Part § 90.259(b)(4) where WMTS may operate in the 1429–1431.5 MHz band.

#### General Technical Requirements

This section only describes general technical requirements. The specific requirements for wireless medical telemetry devices operating in the 608–614 MHz band may be obtained in CFR Part § 95.1119. The specific requirements for wireless medical telemetry devices operating in the 1395–1400 MHz and 1427–1429.5 MHz bands may be referred to CFR Part § 95.1121.

(a) Field strength limits.

- In the 608–614 MHz band, the maximum allowable field strength is 200 mV/m as measured at a distance of 3 meters using measuring instrumentation with a CISPR (The Comité international spécial des perturbations radioélectriques) quasi-peak detector.

- In the 1395–1400 MHz and 1427–1429.5 MHz bands, the maximum allowable field strength is 740 mV/m, as measured at a distance of 3 meters, using measuring equipment with an averaging detector and a 1 MHz measurement bandwidth

(b) Undesired emissions.

- Out-of-band emissions below 960 MHz are limited to 200 microvolts/meter, as measured at a distance of 3 meters, using measuring instrumentation with a CISPR quasi-peak detector.

- Out-of-band emissions above 960 MHz are limited to 500 microvolts/meter as measured at a distance of 3 meters, using measuring equipment with an averaging detector and a 1 MHz measurement bandwidth.

(c) Emission types.

A wireless medical telemetry device may transmit any emission type appropriate for communications in this service, except for video and voice. Waveforms such as electrocardiograms (ECGs) are not considered video.

(d) Channel use.

- In the 1395–1400 MHz and 1427–1429.5 MHz bands, no specific channels are specified. Wireless medical telemetry devices may operate on any channel within the bands authorized for wireless medical telemetry use in this part.

- In the 608–614 MHz band, wireless medical telemetry devices utilizing broadband technologies such as spread spectrum shall be capable of operating within one or more of the following channels of 1.5 MHz each, up to a maximum of 6 MHz, and shall operate on the minimum number of channels necessary to avoid harmful interference to any other wireless medical telemetry devices.

  - 608.0–609.5 MHz

  - 609.5–611.0 MHz

  - 611.0–612.5 MHz

  - 612.5–614.0 MHz

- Channel usage is on a co-primary shared basis only, and channels will not be assigned for the exclusive use of any entity.

- Authorized health care providers, in conjunction with the equipment manufacturers, must cooperate in the selection and use of frequencies in order to reduce the potential for interference with other wireless medical telemetry devices, or other co-primary users. Operations in the 608–614 MHz band (television channel 37) are not protected from adjacent band interference from broadcast television operating on channels 36 and 38.

(e) Frequency stability.

Manufacturers of wireless medical telemetry devices are responsible for ensuring frequency stability such that an emission is maintained within the band of operation under all of the manufacturer’s specified conditions.

#### Radiation Testing

In addition to the technical standards with which all WMTS transmitters must comply to ensure interference-free operation, all WMTS transmitters that operate within 20 centimeters of a person's body must be routinely evaluated to demonstrate compliance with the FCC's radiofrequency (RF) radiation exposure guidelines. However, WMTS transmitters that operate at 20 centimeters or more from a person's body are not required to undergo routine evaluation to demonstrate RF exposure compliance. Portable WMTS transmitters are those designed to be used with a separation distance of less than 20 centimeters from the radiating structures of a device and a person's body.

Mobile WMTS transmitters are those designed to normally operate with a separation distance of at least 20 centimeters from the radiating structures of a device and a person's body. Since medical telemetry devices are typically worn on the body of the patient, it is expected that most WMTS transmitters will be classified as portable transmitters and will therefore be required to demonstrate RF exposure compliance with respect to the SAR (Specific Absorption Rate) limit (specified in 47 CFR 2.1093) Nonetheless, there may be situations in which the WMTS transmitter is mounted on a patient bed or incorporated within a separate device that is more than 20 centimeters away from the patient and nearby persons.

In those cases, the WMTS transmitter will not be required to undergo routine evaluation to demonstrate RF exposure compliance because, as a consequence of the relatively low power at which the transmitters must operate (i.e., less than 1.5 watts effective radiated power) and the separation distance between the transmitter and the body of persons, the potential for mobile WMTS devices to exceed RF exposure limits is substantially remote.

### Medical Device Radiocommunications Service

Prior to the establishment of the MedRadio Service (originally Medical Implant Communications Service (MICS)) in 1999, medical implant devices had to be magnetically coupled to external programmers or readers. This magnetic coupling required that the device implanted in the patient be in very close proximity to the external monitoring/control equipment, often necessitating body contact for proper operation. In addition, medical implant devices operated with very slow data rates, sometimes requiring up to fifteen minutes for the required data transfer. The FCC established the MICS to overcome these limitations of medical implant devices. The FCC concluded, with the agreement of representatives of the medical community and equipment manufacturers, that establishing a MICS would greatly improve the utility of medical implant devices by allowing physicians to establish high-speed, easy-to-use, reliable, short-range (six feet) wireless links to connect such devices with monitoring and control equipment.

MedRadio permits physicians and their patients to take advantage of the benefits of wireless technology to improve the medical care and capabilities of implanted medical devices, thereby improving these patients' quality of life. MedRadio not only permits faster data transfer rates between medical implant devices and external monitoring/control equipment, it reduces the risk of infection to patients, enhances the comfort of patients, and expands the freedom of movement of medical personnel working with the equipment. The MICS rules took effect on January 14, 2000. The rules changing the name of the service to MedRadio and expanding the designated frequency band took effect on August 12, 2009.

In summary, the FCC has established the new MedRadio Service in Part 95 of its Rules, and has incorporated into this new radio service the rules that currently govern the Medical Implant Communications Service (MICS) at 402-405 MHz.  It also allocated to the MedRadio Service two new megahertz of spectrum in adjacent “wing” bands at 401-402 MHz and 405-406 MHz.  Typical MICS operations involve a medical implant device that includes a radiofrequency transmitter operating in conjunction with an external programmer/control transmitter. The rules for MICS did not accommodate body-worn devices, and all such devices had to incorporate LBT technologies.

Additional Spectrum

In addition to 402-405 MHz, devices may be operated in the 401-402 MHz and 405-406 MHz bands.  There are slight differences in the types of devices that may be accommodated in each band segment.

Body-Worn Devices

These devices are defined as having an antenna on the patient-worn device that is placed upon or within a few centimeters of the body.  Body-worn devices may operate only at 401-402 or 405-406 MHz.  However, in order to allow physicians to evaluate the efficacy of a proposed treatment involving implanted devices, temporary body-worn devices may also operate in the 402-405 MHz band subject to certain time limits and power limits.

Eligibility and Permissible Communications MedRadio Service devices may only be used for non-voice communications for diagnostic and therapeutic purposes, and only when the device is provided to a human patient by an authorized health care professional.  The FCC specifically declined to extend authorization to applications involving non-human test subjects in a research laboratory.

#### Summary of MedRadio operations rules

Rules applicable to MedRadio specify that [6]:

- MedRadio transmitters that are not excepted from the frequency monitoring requirements [7] may not operate with an effective isotropic radiated power (EIRP) greater than 25 microwatts. MedRadio transmitters that are excepted from the frequency monitoring requirements may not operate with an EIRP greater than 100 nanowatts in the 402-405 MHz band, 250 nanowatts in the 401-401.85 MHz and 405-406 MHz bands, and 25 nanowatts in the 401.85-402 MHz band.

- MedRadio transmitter emissions are limited to an authorized bandwidth of 300 kHz in the 402-405 MHz band, 100 kHz in the 401-401.85 MHz and 405-406 MHz bands, and 150 kHz in the 401.85-402 MHz band.

- For frequency monitoring, the FCC retained the current requirement for devices at 402-405 MHz to monitor for other transmissions and to select an unoccupied channel on which to operate (i.e., LBT).  However, the FCC will allow non-LBT access methods at 401-402 and 405-406 MHz for both implant and body-worn devices, subject to a power limit of 250 nanowatts EIRP, a maximum duty cycle limit of 0.1 percent and a maximum of 100 communication sessions per hour.  In order to conform with international standards, non-LBT methods may also be used for implant devices within a maximum 300 kHz bandwidth channel centered at 403.65 MHz, with a maximum EIRP of 100 nanowatts, and maximum duty-cycle and session limits of 0.01 percent and 10 per hour, respectively.  To accommodate some existing body-worn devices, non-LBT devices may operate up to 25 microwatts EIRP at 401.85-402 MHz.

- MedRadio transmitters may transmit any emission type appropriate for communications in this service, but may not be used for voice communications.

- MedRadio transmitter out-of-band emissions, as well as in-band emissions that are more than 150 kHz away from the intended center frequency, must be attenuated to FCC-prescribed levels.

- MedRadio transmitters must be tested for emissions and EIRP limit compliance in accordance with FCC-specified procedures.

- MedRadio transmitters must maintain a frequency stability of +/-100 ppm of the operating frequency.

#### Equipment and Interference Safeguards

MedRadio systems consist of the transmitters connected to medical implant or body-worn devices, and programming, monitoring and control equipment. A MedRadio device is generally regarded as apparatus that is placed inside the human body, or placed upon or in very close proximity to the body, for the purpose of performing diagnostic or therapeutic functions, such as cardiac pacemakers and defibrillators.

Users of MedRadio transmitters must cooperate in the selection and use of channels in order to reduce interference and to make the most effective use of the authorized facilities. Most importantly, channels must be selected so as to avoid interference to other MICS transmissions. As a safeguard against such MedRadio-to-MedRadio interference, external medical implant programmer/control MedRadio transmitters operating in the 402-405 MHz band must incorporate a mechanism for monitoring the channel or channels that the MedRadio system devices intend to occupy and, unless there is a “medical implant event” may not initiate a MedRadio communications session unless certain "access criteria" are met [7].

A medical implant event is defined as an occurrence or the lack of an occurrence recognized by a medical implant device, or a duly authorized health care professional, that requires the transmission of data from a medical implant transmitter in order to protect the safety or well-being of the person in whom the medical implant transmitter has been implanted. The Commission has clarified that regularly scheduled transmissions that are not instigated by a change in the patient’s medical condition do not qualify as medical implant events. A request for waiver to permit certification of implanted devices that emit periodic scheduled transmissions is pending.

In addition, all MedRadio operations, as a consequence of their secondary status, must not cause harmful interference to stations in the Meteorological Aids, Meteorological Satellite, or Earth Exploration Satellite Services operating in the same or adjacent frequencies. Further, MedRadio stations must accept interference from such Meteorological Aids, Meteorological Satellite, or Earth Exploration Satellite Service stations.

#### Radiation Testing

MedRadio transmitters must be evaluated by manufacturers to ensure that their use will not result in the exposure of either patients or medical professionals to radiofrequency (RF) radiation emissions that are considered unsafe. Manufacturers are required to file with the FCC, among other things, a certification of compliance with the RF exposure rules in order to secure FCC approval to market their equipment.

## UWB Frequency Band

The unlicensed UWB operation regulations are described in subpart F of part 15 of 47 CFR.

### Definitions

(a) UWB Bandwidth: The UWB bandwidth is the frequency band bounded by the points that are 10 dB below the highest radiated emission, as based on the complete transmission system including the antenna. The upper boundary is designated by FH and the lower boundary is designated by FL. The frequency at which the highest radiated emission occurs is designated by FM .

(b) Center frequency: The center frequency (FC) equals to (FH + FL)/2.

(c) Fractional bandwidth: The fractional bandwidth equals to 2(FH - FL)/ (FH + FL).

(d) Ultra-wideband (UWB) transmitter: An intentional radiator that, at any point in time, has a fractional bandwidth equal to or greater than 0.20 or has a UWB bandwidth equal to or greater than 500 MHz, regardless of the fractional bandwidth.

(e) EIRP: Equivalent isotropic radiated power.

(f) Hand held: A hand held device is a portable device, such as a lap top computer or a PDA that is primarily hand held while being operated and that does not employ a fixed infrastructure.

(g) Medical imaging system: A field disturbance sensor that is designed to detect the location or movement of objects within the body of a person or animal.

(h) Indoor UWB devices, by the nature of their design, must be capable of operation only indoors. The necessity to operate with a fixed indoor infrastructure, e.g., a transmitter that must be connected to the AC power lines, may be considered sufficient to demonstrate this.

### General technical requirements

- UWB devices may not be employed for the operation of toys.

- Operation onboard an aircraft, a ship or a satellite is prohibited.

- Emissions from digital circuitry used to enable the operation of the UWB transmitter shall not exceed the field strength level of 500 microvolts/meter at measured distance of 3 meters for 960 MHz and above. A tighter emission limit applies at the band edges for some systems. The EIRP (in dBm) values versus frequency ranges are listed in the following table when measured using a resolution bandwidth of 1MHz.

Table 9.1: EIRP (in dBm) versus Frequency range

| frequency range (MHz) |  | medical imaging systems | hand held UWB systems | Indoor UWB systems |
| --- | --- | --- | --- | --- |
| 960 | 1164 | -65.3 | -75.3 | -75.3 |
| 1164 | 1240 | -65.3* | -75.3** | -75.3** |
| 1240 | 1559 | -65.3 | -75.3 | -75.3 |
| 1559 | 1610 | -65.3* | -75.3** | -75.3** |
| 1610 | 1990 | -53.3 | -63.3 | -53.3 |
| 1990 | 3100 | -51.3 | -61.3 | -51.3 |
| 3100 | 10600 | -41.3 | -41.3 | -41.3 |
| 10600 | above | -51.3 | -61.3 | -51.3 |

*: The EIRP limit is -75.3dBm when measured using a resolution bandwidth of no less than 1 kHz

**: The EIRP limit is -85.3dBm when measured using a resolution bandwidth of no less than 1 kHz

- Radiated emission levels [8] at and below 960 MHz are based on measurements employing a CISPR quasi-peak detector.

- Radiated emission levels above 960 MHz are based on RMS average measurements over a 1 MHz resolution bandwidth. The RMS average measurement is based on the use of a spectrum analyzer with a resolution bandwidth of 1 MHz, an RMS detector, and a 1 millisecond or less averaging time. Unless otherwise stated, if pulse gating is employed where the transmitter is quiescent for intervals that are long compared to the nominal pulse repetition interval, measurements shall be made with the pulse train gated on. Alternative measurement procedures may be considered by the Commission. [The reference to a 1 millisecond or less averaging time denotes the integration time period for each bin on the spectrum analyzer. The analyzer sweep time/number of bins must not exceed one millisecond. Trace averaging should not be used in making the RMS measurements.]

- The frequency at which the highest radiated emission occurs, FM , must be contained within the UWB bandwidth.

- Imaging systems may be employed only for the type of information exchange described in their specific definitions contained in previous section. The detection of tags or the transfer or data or voice information is not permitted under the standards for imaging systems.

When a peak measurement is required, it is acceptable to use a resolution bandwidth other than the 50 MHz specified in this subpart. This resolution bandwidth shall not be lower than 1 MHz or greater than 50 MHz, and the measurement shall be centered on the frequency at which the highest radiated emission occurs, FM . If a resolution bandwidth other than 50 MHz is employed, the peak EIRP limit shall be 20 log (RBW/50) dBm where RBW is the resolution bandwidth in megahertz that is employed. This may be converted to a peak field strength level at 3 meters using E(dBuV/m) = P(dBm EIRP) + 95.2.

# ISM Band

This section presents a short description on the Industrial, Scientific, and medical (ISM) frequency bands. The ISM bands are allocations for unlicensed operation, thus are potential candidates to be considered by 15.6 TG.

The ISM bands are defined by the ITU-R in 5.138, 5.150, and 5.280 of the radio regulations. Individual countries' use of the bands designated in these sections may differ due to variations in national radio regulations. Because communication devices using the ISM bands must tolerate any interference from ISM equipment, these bands are typically given over to uses intended for unlicensed operation, since unlicensed operation typically needs to be tolerant of interference from other devices anyway. ITU-R defines a large rang of frequency bands for ISM as follows (Some of them are subject to local acceptance):

- 6.765–6.795 MHz (centre frequency 6.780 MHz)

- 13.553–13.567 MHz (centre frequency 13.560 MHz)

- 26.957–27.283 MHz (centre frequency 27.120 MHz)

- 40.66–40.70 MHz (centre frequency 40.68 MHz)

- 433.05–434.79 MHz (centre frequency 433.92 MHz)

- 902–928 MHz (centre frequency 915 MHz)

- 2.400–2.500 GHz (centre frequency 2.450 GHz)

- 5.725–5.875 GHz (centre frequency 5.800 GHz)

- 24–24.25 GHz (centre frequency 24.125 GHz)

- 61–61.5 GHz (centre frequency 61.25 GHz)

- 122–123 GHz (centre frequency 122.5 GHz)

- 244–246 GHz (centre frequency 245 GHz)

Among the above frequency bands, some are regional. For example,

- 902 -928 MHz band

This band is governed by Part 18 of FCC of USA and then it is also used by many regulations in other regions or countries. This band is allowed to be used without license in North America, Australia, New Zealand, and a number of other nations.

In addition, there are also regional available bands such as

- 868.0 – 870.0 MHz band

This band is recommended by the European Telecommunications Standards Institute (ETSI) and is allowed to be used without license in most EU countries.

It should be noticed that some world wide available ISM bands have been used by many existing standards. For example, a number of IEEE standards use the 2.4 – 2.4835 GHz band including 802.11b. g, (wireless LAN), 802.15.1 (Bluetooth), 802.15.4 (Zigbee), and 802.15.4a (CSS PHY), etc. Coexistence issue must be paid enough attentions for such band. Moreover, besides the frequency regulation on ISM band, there are also regulations for unlicensed communication devices. Those rules for unlicensed communication devices must be taken into consideration even for ISM band.

# Related regulation for HBC

This section presents information on HBC related frequency regulations according to reference [9].

## EU

- Regulation of Short Range Devices (SRDs) [ETSI EN 300 330]

  - SRD is defined as the radio transmitter which provide either unidirectional or bi-directional communication and which have low capability of causing interference to other radio equipment

  - For SRDs, individual licenses are normally not required

  - No restricted bands unlike USA, Korea, and Japan

## Japan

Regulation of the Extremely Low Power Radio Station is as follows

- When the level of the electric field intensity (the strength of the radio wave) within 3 meters is below 500μV/m (fc< 322MHz), there is no need for getting a license from the radio station.

- Japan Radio Law Operation Regulation –Article 6, Section 1

- Japan Radio Law Article 4, Section 1

## Korea

- Regulation of the Extremely Low Power Radio Station (ELPRS)

  - When the level of the electric field intensity within 3 meters is below 500μV/m (fc< 322MHz), there is no need for getting a license from the radio station.

  - Same as Japan

- Radio Station Law

  - Article 97

  - Except in the restricted bandslisted in the following table

Restricted frequency band (10 – 50 MHz)

## UK

- Licence Exempt Short Range Devices

  - UK Interface Requirement 2030

  - The UK Interface Requirement defines the radiation level for the use of Licence Exempt Short Range Devices in the specified bands

  - No restricted bands

  - In the case of general non-specific short range devices (section 3)

## USA

### FCC Regulations for low-power, non-licensed transmitters

- If a particular transmitter can comply with the general radiated limits (Section 15.209), and at the same time avoid operating in one of the restricted bands (Section 15.205), then it can use any type of modulation (AM, FM, PCM, etc.) for any purpose.

- Section 15.209 contains general radiated emission (signal strength) limits that apply to all Part 15 transmitters using frequencies at 9 kHz and above. (see table 11.1)

Table 11.1 Emission limits for frequency of 9 kHz and above

- Section 15.205 lists a number of restricted bands in which low power, non-licensed transmitters are not allowed to operate.

### FCC Regulations for low-power, non-licensed transmitters

Table 11.2 summarizes the restricted frequency band of section 15.205.

Table 11.2 Restricted frequency band (section 15.205)

# Summary

Frequency bands described are shortly summarized in Figure 12.1.

Figure. 12.1 Frequency regulation summary

(In EU and USA, 401-402 MHz and 405-406 MHz can also be used for MICS)

- MICS band is widely accepted although the possible bandwidth is limited.

- WMTS band has different definitions in different countries or regions.

- Some ISM bands are available world wide while coexistence may be an issue.

- UWB bands are available world wide while some spectrums are defined differently in different countries or regions.

In order to give a concise summary on different frequency bands, brief views on MICS, WMTS, UWB low band, as well as UWB high band are given in Table 12.1 through 12.4.

Table 12.1 A brief view on MICS band

*Readers are also encouraged to refer to ITU Recommendation RS.1346 and ETSI EN301839-1, -2.

Table 12.2 A brief view on WMTS band

Table 12.3 A brief view on UWB low band

Table 12.4 A brief view on UWB high band