IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs) |  |
| --- | --- | --- |
| Title | TG4j Draft 3 Annex P1 Coordinator Switch revised text |  |
| Date Submitted | [9 August 2012] |  |
| Source | [Dave Evans] [Philips] [101 Cambridge Science Park, Cambridge, UK] | Voice: [+44 1293 886490] Fax: [ ] E-mail: [david.evans@philips.com] |
| Re: | [This is a proposed revision for TG4j Draft 2.] [If this is a response to a Call for Contributions, cite the name and date of the Call for Contributions to which this document responds, as well as the relevant item number in the Call for Contributions.] [Note: Contributions that are not responsive to this section of the template, and contributions which do not address the topic under which they are submitted, may be refused or consigned to the “General Contributions” area.] |  |
| Abstract | [Revised text for Annex P1 in response to Letter Ballot 82 to address the proposed changes from the three editorial comments and new text to address the one editorial comment on this annex.] |  |
| Purpose | [Description of what the author wants P802.15 to do with the information in the document.] |  |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |  |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |  |

Annex P

(informative)

Features to assist MBAN devices

This annex describes two features that could be used to assist the operation of MBAN devices.

P.1 Coordinator switching

There are several MBAN usage scenarios where the devices that are associated with one coordinator may need to be moved to another coordinator. For example a patient with an array of sensors is in transportation from an operating room to a recovery room and monitored by a portable monitor to provide continuous care services. The patient is then moved back to his/her patient room in which a bedside monitor is available. Switching the sensor devices from the portable monitor to bedside monitor would provide better data processing and display and save the battery life of the portable device. It is important that an orderly switching of the devices from the coordinator in the portable monitor to the desired coordinator in the bedside monitor takes place. Since the switching is instigated and controlled by a clinician and the procedure to enable this resides in the higher layer of the coordinators, a detailed description of how this is done is beyond the scope of this standard though some of the steps that could be used are noted here.

Firstly the coordinator sends a message to the associated devices informing the devices of the identity of the coordinator that they should associate with after they have been disassociated from their existing coordinator. This identity information is supplied to the coordinator by the clinician. On receipt of this information the devices should then cease further data transmission and await notification from the coordinator that they have been disassociated. Upon disassociation, the devices then commence a scanning procedure to find the new coordinator whose identity is being broadcast within the beacon to determine the channel number that the coordinator is using, its PAN ID and MAC extended address. Once this information is obtained the devices can begin association and then communication between the devices and the new coordinator can commence. This description does not necessarily include all of the steps that need to be taken and it only gives an indication of what may be required.

P.2 Channel bitmap

To assist with coexistence to the primary users of the 2360-2390 MHz band, MBAN devices may be excluded from using some of the channels in this band (channels 0-5 and 7-12). The channels in the 2390-2400 MHz band (channels 6, 13 and 14) are always available for use. The channels that are excluded in the 2360-2390 MHz band may change over time. Regulations require that the operation of MBAN hub devices have access to a mechanism that provides information on the portions of the band that cannot be used. The description of this mechanism is beyond the scope of this standard. To assist other MBAN devices, such as those connected to sensors, that do not have direct access to this mechanism, a bitmap may be used by a hub device to indicate the channels that are allowed to be used by the MBAN devices.

The bitmap can be conveyed as an Encapsulated Service Data Unit (ESDU) Information Element (IE) described in 5.2.4.4 that can be used to encapsulate a higher layer payload. The definition of the format of this payload is beyond of the scope of this standard.