## Slide 1
Project: IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs)
Submission Title:[ Review of IEEE802.15.6 Wireless Medical BAN]
Date Submitted: [12 July 2018]
Source: [Ryuji Kohno1,2,3, Huan-Bang Li 4 ] [1;Yokohama National University, 2;Centre for Wireless Communications(CWC), University of Oulu, 3;University of Oulu Research Institute Japan CWC-Nippon
4; National Institute of Information and Communications Technology(NICT)]
Address [1; 79-5 Tokiwadai, Hodogaya-ku, Yokohama, Japan 240-8501
2; Linnanmaa, P.O. Box 4500, FIN-90570 Oulu, Finland FI-90014
3; Yokohama Mitsui Bldg. 15F, 1-1-2 Takashima, Nishi-ku,Yokohama, Japan 220-0011
4: 3-4, Hikarino-oka, Yokosuka-city, Japan 239-0847]
Voice:[1; +81-45-339-4115, 2:+358-8-553-2849], FAX: [+81-45-338-1157],
Email:[1: kohno@ynu.ac.jp, 2: Ryuji.Kohno@oulu.fi, 3: ryuji.kohno@cwc-nippon.co.jp, 4:lee@nict.go.jp]
Re: []
Abstract: [In order to reconsider IG-DEP, first its activities is reviewed and As a dependable MAC protocol for wireless body area network(WBAN), a scheme of media access control specified for bi-directional transmission of related packets between a coordinator and nodes of is introduced.]
Purpose: [information]
Notice: This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein.
Release: The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15.

## Slide 2: Review of IEEE802.15.6 Wireless Medical BAN
Ryuji Kohno1)2)3), Huan-Bang Li4)
1) Yokohama National University, Japan
2) University of Oulu, Finland
3) University of Oulu Research Institute Japan – CWC-Nippon, Co. Ltd.
4) National Institute of Information and Communication Technology, Japan

## Slide 3: Agenda
Short Summary of IG-DEP Activities
Review of Previous Meetings
Definition of Dependability in Wireless Networks
Use Cases and Applications of Dependable Wireless
Focused Use case; Automotive Use Cases
Summary of Technical Requirement for Automotive
Technical Requirement for Automotive and UAV
Demand of Medical Infrastructure Platform of BAN, Cloud Networks and AI Data Mining Server and Repository
Review of IEEE802.15.6 for Medical WBAN
Discussion; Possible Amendment of IEEE802.15.6

## Slide 4: 1. Wireless BAN: Body Area Network
Tele-metering or sensing
vital signs with various sensors
Wearable BAN
Implant BAN
Tele-control of Medical
Equipment and Devices
ECG
EEG
Blood Pressure
Heart Beat
Body temperature
Sugar rate
Medical images
And video
Etc.
Pace Maker with ICD
Novel Concept
Intelligent Network of Vital
Sensors, eHR, Medical Robots etc.
Wireless Capsule Endoscope
Notes: 4

## Slide 5
2. Standard of Medical Wireless Body Area Network(BAN);IEEE802.15.6
802.15.4b
Revision & Modification of 15.4 MAC
802.15.3
PHY for High Rate WPAN
IEEE802.11
Wireless LAN
802.15.2
Coexistence between WPAN and WLAN
802.15.5
WPAN
Mesh Network
IEEE802.15
Wireless PAN
802.15.1
Bluetooth
802.15.4
PHY for Low Rate WPAN
802.15.3a
Alternative PHY of 15.3
802.15.3c
PHY in Millimeter wave band
802.15.3b
Maintenance of 15.3
802.15.4a
Low rate UWB PAN Alternative PHY of 15.4
IEEE802
IEEE802.16
Wireless MAN
802.15.4c
Chinese
WPAN
802.15.4d
Japanese
WPAN
2007.3
Standard Completed
802.15.6
Wireless
Medical BAN
2012.2 Standard
was Completed

## Slide 6
3. Main Contributors at TG6
Casuh
CEA-LETI
CNU
CSEM
CUNY
ETRI
France Telecom
Fujitsu Lab. Europe
Fujitsu Lab.
GE Global Research
GE Healthcare
IMEC
Inha University
KETI
Korpa
LG Electronics
Meiji Univesity
Mitsubishi Electric Research Labs, USA
NICT
NICTA
NIST
Olympus, USA
Philips, USA
Philips, EU
Samsung
Tensorcom
Texas Instrument
Thales
Toumaz Technologies
Yokohama National University
Zarlink Semiconductor
… …
USA
Europe
Asia

## Slide 7
4. Overview of IEEE Std 802.15.6
UWB: Ultra-wideband
HBC: Human body communication
IEEE 802.15.6
Narrow band PHY
on-body & in-body
UWB PHY
on-body
Common MAC
（for all PHY）
Coexistence?
Outage probability?
Power consumption?
Complexity?
Security?
Reliability?
Modulation: GMSK
& DPSK
TX range: ~3m
Bands: MICS,WMTS,ISM
Data rate: ~ some Mbps
Modulation:IR-UWB
& FM-UWB
TX range: ~3m
Band: UWB band
Data rate: ~10Mbps
Beacon-base-TDMA
Group Superframe
Priority support
Non-beacon mode
HBC PHY
on-body
Frequency Selective
10-50MHz
125kbps-2Mbps

## Slide 8: 5. 15.6 User Priority Mapping
Priority level | Traffic designation | Data type
7 | Emergency or medical event report | Data
6 | High priority medical data or network control | Data or management
5 | Medical data or network control | Data or management
4 | Voice | Data
3 | Video | Data
2 | Excellent effort | Data
1 | Best effort | Data
0 | Background | Data

## Slide 9: 6. Three Channel Access Modes
Channel access mode | Time reference-based (superframe structure) | Beacon | Notes
I | Yes | Yes | Coordinator sends beacon in each superframe except for inactive superframes.
II | Yes | No | Coordinator establishes time reference but doesn’t send beacon.
III | No | No | There is not time reference.

## Slide 10
7. Time-referenced Superframe w/ Beacon
EAP: exclusive access phase
RAP: random access phase
MAP: managed access phase
CAP: contention access phase
Clock and position of each access phase
May obtain contended allocation for highest priority
B
B2
EAP1
RAP1
MAP
EAP2
RAP2
MAP
CAP
One superframe

## Slide 11
8. Main Features of the Three PHYs
 | Frequency band (MHz) | Data rate (kbps) | Note
NW-PHY | 400, 600, 800, 900, 2400 | 75.9 --- 971.4 | Interference with other systems operate at the same bands
UWB-PHY | 6000-10600 3100-4800 | 390 --- 12600 | Worldwide common band is 7.25 – 8.5 GHz
HBC-PHY | 21 | 164 --- 1312.5 | Strong concern on the effect to implant devices

## Slide 12
9. Main Specifications of NB-PHY
Frequency bands (MHz) | Modulations |  | Data rates (kbps) | Number of channel | Notes
 | PLCP header | PSDU |  |  | 
402-405 | π/2- DBPSK | π/2-DBPSK, π/4-DQPSK π/8-D8PSK | 75.9/151.8/ 303.6/455.4 | 10 | Majority of countries
420-450 | GMSK | GMSK | 75.9/151.8/187.5 | 12 | Japan
863-870 | π/2-DBPSK | π/2-DBPSK, π/4-DQPSK π/8-D8PSK | 101.2/202.4/ 404.8/607.1 | 14 | EU
902-928 | π/2-DBPSK | π/2-DBPSK, π/4-DQPSK π/8-D8PSK | 101.2/202.4/ 404.8/607.1 | 60 | North America, Australia
950-958 | π/2-DBPSK | π/2-DBPSK, π/4-DQPSK π/8-D8PSK | 101.2/202.4/ 404.8/607.1 | 16 | Japan
2360-2400 | π/2-DBPSK | π/2-DBPSK, π/4-DQPSK | 121.4/242.9/ 485.7/971.4 | 39 | USA
2400-2483.5 | π/2-DBPSK | π/2-DBPSK, π/4-DQPSK | 121.4/242.9/ 485.7/971.4 | 79 | Worldwide

## Slide 13
Mode | Modulation | Data rate (Mbps) | Waveform
IR-UWB (I) | OOK | 0.49 – 15.6 | Chirp pulse, chaotic pulse, SRRC-like pulse, or others.
IR-UWB (II) | DBPSK/DQPSK | 0.49 – 15.6 | 
FM-UWB | Continuous-phase 2FSK (sub carrier) combined with FM | ≦0.25 | Gaussian (default)
FM-UWB is an optional mode
High QoS mode
  Hybrid Type II ARQ
10. Main Specifications of UWB-PHY

## Slide 14
HBC frequency band
  center frequency:21MHz (3dB_BW=5.25MHz）
Transmission method
  Frequency Selective Digital Transmission
Data rate
  164, 328, 656, 1312.5 kbps
11. Main Specifications of HBC PHY
The electrode in contact with the body is used for transmitting or receiving an electrical signal through the body to a device

## Slide 15
12. World Wide UWB Regulations

## Slide 16
13. Specifications of High Band UWB
Items | Specifications
Frequency band | 7.25 – 10.25 GHz
Average e.i.r.p. | ≤ -41.3 dBm/MHz
Peak e.i.r.p. | ≤ 0 dBm/50MHz
Average unwanted radiation | ≤ -70 dBm/MHz
Peak unwanted radiation | ≤ -64 dBm/MHz
Pulse rate | ~ 50 Mpps
Communication range | ~ 3m

## Slide 17: 14. Summary of IEEE802.15.6
  Body area network (BAN) is considered as an core technology in supporting automatic medical monitoring and healthcare agilence services as well as consumer centric electronics.
  A standard, IEEE Std 802.15.6TM was completed and published in Feb. 2012. In which, Unified MAC and three PHY specifications are defined to support different applications.
  Commercial BAN products have been sold for medical healthcare for a human body in a world.
  In PHY, UWB-BAN is appropriate for high QoS use case.
  In MAC, hybrid contention base and free protocols is reliable to guarantee delay and throughput flexibly.

## Slide 18: 15. Demands for BAN Extension
BAN for Car and Other Bodies beyond Human Body
Reliable performance of medical BAN for human body could be widely applicable for remote maintenance of car body and other bodies in IoT/M2M use cases.
Demands for More flexible and widely applicable BAN in cars, robotics, UAVs and others are increasing for autonomous remote sensing and controlling.
current IEEE802.15 IG-Dependability
BAN-base Infrastructure Platform for Medical Healthcare
BANs in end users are connected through Cloud Network and Edge Computer with AI Data Mining Server and Repository for medical healthcare platform by integration between ICT and data science.
Enhanced dependability is required for end-to-end reliability and security.
BAN-base Universal Platform for Medical and beyond Medical Infrastructures
Emergency for natural disasters and terrorism, smart city with reliable maintenance of cars, buildings etc. need common dependable platform,